EV 420 Aligner SOP

Scope

This document provides procedures for operating and maintaining the EV 420 front and backside aligner.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	2

3.1	Referenced within this Document	2

3.2	External Documents	2

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Setup Procedures	3

6.1	Turn on UV Lamp	3

6.2	Replace Mask Holder	3

6.3	Turn on Aligner	4

7.	Front Side Alignment Procedures	4

7.1	Load Mask	4

7.2	Align Mask	4

7.3	Load and Align Sample	5

7.4	Enter Exposure Time	5

7.5	Contact Mask and Sample	5

7.6	Expose and Remove Sample	6

7.7	Remove Mask and Shutdown	6

8.	Back Side Alignment	6

8.1	Align Mask and Sample	6

9.	Process Notes	7

9.1	Exposing Thick SU-8	7

9.2	Adjust Pressures	7

10.	Revision History	7

Figure 1, UV Lamp Power Supply	3

Figure 2, Control Panel	4

Figure 3, Microscope Light Power Supply	5

Reference Documents

Referenced within this Document

None

External Documents

None

Equipment and/or Materials

EV 420 Aligner

Mask

Wafer/Sample

Safety

Follow all Nanofab safety procedures.

Setup Procedures

Turn on UV Lamp

Make sure EV 420 power is off.

Turn on the white power switch on the UV lamp power supply located to the left of the aligner. See Figure 1.

Press ‘start’ on the UV lamp power supply. See Figure 1.

Power Switch

Start Button

Figure 1, UV Lamp Power Supply

Replace Mask Holder

Mask holders are available for both 4” and 5” masks. If the correct mask holder is not on the system do the following:

Remove the microscope objectives

Unscrew each objective

Put one hand underneath the objective as you unscrew them to catch them when they fall

Place the objectives in a safe place where they won’t fall and get damaged.

Remove the mask holder

Unscrew thumbscrews on each side of the mask holder and slide the sliders away

Lift up gently and slide mask holder towards front of the aligner

Put mask holder in the top drawer of the white spinner bench.

Install the new mask holder

Gently slide the mask holder in, with the rim facing the back of the aligner.

Slide in the sliders

Screw in the thumbscrews

Reinstall the microscope objectives.

Turn on Aligner

Press the green Power On button. See Figure 2.

Wait for the system to initialize.

Figure 2, Control Panel

Power On Button

Front Side Alignment Procedures

Load Mask

Move out the tray when prompted.

Load the mask frame with the three plastic pins facing up.

Place the mask face down between the three pins.

Move the tray in.

Wait for the mask to level.

Align Mask

Turn on the microscope light power supply. See Figure 3.

Adjust the voltage until it reads 7.0 V. See Figure 3.

View the mask through the optics.

Rotate the mask and/or the objectives until two similar features appear on both halves and are lined up horizontally.

Make sure that the mask is entirely underneath the white vacuum grooves.

Press Enter when the mask is aligned.

Power/Voltage Knob

Figure 3, Microscope Light Power Supply

Load and Align Sample

Unload the mask frame and press ‘Enter’.

Load your sample and press the ‘V-Chuck’ button.

Press ‘Enter” when it says ‘remove ruler’ on the LCD.

Move the tray in when prompted.

Wait for the system to level. The LCD will read ‘align wafer’ when it is done leveling.

Align the wafer to the mask using the micrometers to rotate the substrate or move it back and forth.

Enter Exposure Time

Press ‘Menu’.

Press the right arrow once (numeric keypad 6) to select ‘Proc’ on the display.

Press ‘Enter’.

Press ‘Enter’.

Press ‘Enter’.

Press ‘Enter’.

Press ‘Enter’.

Enter the desired exposure time.

Press ‘Enter’.

Press ‘Enter’.

Press ‘Menu’.

Contact Mask and Sample

When the substrate and mask are aligned, press the ‘separation’ button.

Press the ‘H-Cont’ button and wait until it says ‘Adjust Chamber Vaccum’.

Make sure the hard-cont pressure gauge reads a max of 1 bar, then press ‘Enter’.

Expose and Remove Sample

Press the ‘EXP’ button when it is selected.

Wait for the prompt, then move tray out and remove sample.

Load another sample if desired (see 7.1).

Remove Mask and Shutdown

Press the red “RESET” key.

Press ‘Enter’ when the LCD reads ‘unload mask’.

Move tray in when prompted.

Move tray out when prompted.

Remove your mask.

Back Side Alignment

Align Mask and Sample

Move tray out when prompted.

Place the mask face down but DO NOT use the mask frame with the three pins.

Move tray in when prompted.

Wait for the mask to level (LCD will read ‘align mask’).

Turn up the backside (bottom) illumination on the control panel.

Press the TB button.

On the TV monitor turn brightness down all the way and turn contrast up all the way.

Turn off the top illumination and the room lights.

Press LR to toggle between moving the microscope objectives.

Press the up and down arrows to focus the left and right objectives.

Align the crosses on the mask to the crosses on the TV screen by rotating and moving the mask left and right with the micrometers.

Press ‘enter’ and the mask vacuum will turn on.

Load the wafer (see 7.1).

Do not touch the bottom focus or move the bottom lenses. Move only the sample when it tells you to align the sample.

Align the crosses on the backside of the sample to the crosses on the TV screen.

Press the ‘Enter’ button when the sample is aligned.

Set the exposure time (see 7.4).

Expose and remove the sample (see 7.6).

Shut down the system (see 7.7).

Process Notes

Exposing Thick SU-8

Use a filter over top of mask

Break up exposure into small time increments (~30 sec. intervals) so SU-8 does not over heat and cure before development.

Menu>Proc>Expo>Exp>Preset>Edit>1

Enter total exposure time in seconds

Int- enter # of breaks

Enter break time (lamp off time)

Proximity off

Enter separation gap (typically 80 um)

Select>1>Menu

Adjust Pressures

Contact Force-Align: 1.1

This is what the left gage usually reads, except during WEC.

Chuck Weight: should be between 0.26-0.29 (bellows under WEC unit)

To calibrate WEC Chuck Weight:

Menu>Proc>Exp>WEC>Cal

Load substrate, V-Chuck, Move tray in

Adjust chuck weight to read 0.241

Controls separation distance between mask and wafer during soft contact and alignment.

Contact Force: 0.5 bar (WEC force)

Adjust the contact force during WEC. This pressure is only displayed during WEC.

The higher the number the lower the contact force.

Hard Contact Pressure: adjust any time (Max 1 bar)

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 12 Jan 2010 | Sam Bell |  |