Oxford 100 Cryo DRIE SOP

Scope

This document provides operating procedures for the Oxford 100 Cryo DRIE.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	2

3.1	Referenced within this Document	2

3.2	External Documents	2

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Setup Procedures	2

6.1	Record Information in Log Book	2

6.2	Turn on Burn Box	2

6.3	Clean the Chamber	2

6.4	Mask Wafer	2

6.5	Turn on the Liquid Nitrogen	2

6.6	Turn on the Chiller	3

6.7	Vent load lock.	3

6.8	Load Wafer	3

6.9	Pump Down Load Lock	3

7.	Cryo Etching Procedure	3

7.1	Transfer Wafer to Etch Chamber	3

7.2	Etching Si with a Photoresist Mask	5

7.3	Control the Etch Profile	5

8.	Shutdown Procedures	6

8.1	Vent Load Lock	6

8.2	Shutdown System	6

9.	Maintenance Procedures	7

9.1	Chamber Cleaning	7

9.2	Check He Leak Rate	7

9.3	Check Chamber Leak Rate	7

9.4	Troubleshooting	7

10.	Process Notes	8

10.1	Typical Etch Characteristics	8

10.2	Process Summary	10

11.	Revision History	10

Figure 1, Chiller	2

Figure 2, Load Lock	3

Figure 3, Pump Controls Page	4

Figure 4, Process Controls Page	4

Figure 5, Etch Profile	6

Table 1, Cleaning Recipe Parameters	7

Table 2, Cryo DRIE Parameters	10

Reference Documents

Referenced within this Document

Ecosys Burn Box SOP

External Documents

None

Equipment and/or Materials

Wafer/Sample

Oxford 100

Liquid Nitrogen

Safety

Follow all Nanofab safety procedures.

Include all other safety precautions necessary for these procedures.

Setup Procedures

Record Information in Log Book

Record all processing and characterization information in log book.

Turn on Burn Box

Follow the procedure in the Ecosys Burn Box SOP to turn on the burn box.

Clean the Chamber

Run the cleaning recipe as described in for 25 minutes for every hour of Bosch use. See Section 9.1 Chamber Cleaning.

Power Switch

Figure 1, Chiller

Mask Wafer

For etch depths less than 20 um, S1813 spun at 3000 RPM can be used. For depths greater than 20 um, it is highly suggested that you use a thicker resist. The etch selectivity of Si:S1813 is about 10:1.

Turn on the Liquid Nitrogen

Open the liquid nitrogen valve.

Turn on the Chiller

Go into the chase behind the Oxford 100.

Turn the power switch to start on the chiller. See Figure 1, Chiller.

Vent load lock.

Click on system icon. See Figure 3, Pump Controls Page.

Click on pump page.

Press stop button corresponding to the load lock mechanical pump. See Figure 3, Pump Controls Page

Press vent button. See Figure 3.

Wait for load lock to vent. The pressure will read above 600 Torr and the time will be 0.

Figure 2, Load Lock

Load Wafer

Open load lock lid.

Press stop button.

Place wafer in load lock against the two pins on the transfer arm with the wafer facing the two pins. The wafer should touch the two pins. See Figure 2, Load Lock.

Pump Down Load Lock

Close load lock lid.

Select the evacuate button. See Figure 3.

Enter an ID name or number for your wafer.

Wait for the load lock to pump down (~1 min).

Cryo Etching Procedure

NOTE:	This recipe can be used to etch smooth shallow trenches in 100 mm wafers. It is strongly advised that you do a test run on a practice wafer before working with your device wafer.

Transfer Wafer to Etch Chamber

Click on the icon for the etch chamber. See Figure 3.

Click on the icon for the load lock chamber.

Click on the etch chamber icon.

System Button

Process Buttton

Load Lock Controls

Evacuate

Stop

Vent

Main Chamber Controls Do Not Touch

Figure 3, Pump Controls Page

Load Lock Icon

Main Chamber Icon

Figure 4, Process Controls Page

Start Button

RF Power

ICP Power

He Backing Pressure

Temperature

Etching Si with a Photoresist Mask

Click Process button. See Figure 3, Pump Controls Page.

Select Chamber 1.

Verify that the “ignore tolerances” box is checked.

The temperature should be -120 C. See Figure 4, Process Controls Page.

Allow the chamber to pump down to ~1x10^-6 T (approx. 15 min).

Set the run time to 20 min.

Set the He backing pressure to 10mT.

Press the Start button. See Figure 4.

Allow He to flow for 3 min.

This should give a chamber pressure of 1-5x10^-5T.

Set the SF6 to 34.5sccm and the O2 to 5.5sccm.

Press the Start button.

The pressure should now be 3-6mT.

The pressure set point should be 0mT.

Set your desired etch time.

Set the RF power to 35W and the ICP power to 500W.

Click the Start button to begin your process.

Once the plasma ignites, immediately change the RF power to 5W.

Click Start button.

If plasma does not strike, stop the process and repeat steps with a higher striking RF power.

Control the Etch Profile

The etch profile can be controlled by adjusting the SF6 flow rate in the above recipe.

Increasing the SF6 will create a negative profile while doing the opposite creates a positive profile. See Figure 5.

a)

b)

c)

d)

e)

f)

Change in etch profile as a function of SF6. a) SF6=37, b) SF6=36, c) SF6=35, d) SF6=34, e) SF6= 33, and f) SF6=32 sccm.

Figure 5, Etch Profile

Shutdown Procedures

Vent Load Lock

Click OK when the software says “Process completed”.

Vent load lock.

Press stop button corresponding to the load lock mechanical pump.

Click ok when it says “Wafer has finished processing”.

Press vent button. See Figure 3.

Wait 3 min. for the load lock to vent.

Pull up to open load lock lid. Do not force it open, once it is vented it should open easily. If it does not open easily, press stop, then vent, and wait 3 minutes again.

Press stop button to stop venting.

Remove wafer.

Shutdown System

Close valve on liquid nitrogen tank.

Turn off the chiller.

Pump down load lock.

Close load lock lid.

Select the evacuate button.

Press ‘cancel’ when software asks for ID #.

Turn off burn box.

Follow procedures in the Ecosys Burn Box SOP to turn off the burn box.

Maintenance Procedures

Chamber Cleaning

Chamber cleaning should be done after every 3-5 hours of etching or as needed.

The cleaning recipe parameters are as follows:

| Table 1, Cleaning Recipe Parameters |  |
| --- | --- |
| - SF6 | - 10sccm |
| - O2 | - 40sccm |
| - ICP power | - 1000W |
| - RF power | - 100W |
| - He back pressure | - 0mT |
| - Pressure | - 20mT |
| - Temperature | - Room Temp. |

Check He Leak Rate

Check chamber pressure.

Check icon to make sure wafer is clamped.

Set He pressure to 10 mT (pressure greater than 20 Torr can damage thin substrates).

Check He flow rate. It should be less than ?? sccm.

Check chamber pressure, it should have risen less than 7 mTorr.

Check Chamber Leak Rate

Click on Chamber 1.

Click on leak detection.

Pump down over night.

Run leak test (step time) for 20 minutes.

Log interval 5 sec.

Pumping to 7e-7 Torr

Click return to process.

Troubleshooting

To remove clamped wafer when the system doesn’t recognize there is a wafer in the chamber:

Pump down load lock.

Click on service mode.

Click on location where the wafer is located in the chamber.

Click “add wafer”.

Exit service mode.

Click on green wafer and blue arrow should appear.

Pump down load lock.

Click on center of load lock to transfer wafer to load lock.

Process Notes

Typical Etch Characteristics

The etch rate for aspect ratios less than 5:1 is approx. 1.75 um/min. For higher aspect ratios the etch rate decreases.

Process Summary

The Cryo DRIE parameters are as follows:

| Table 2, Cryo DRIE Parameters |  |
| --- | --- |
| - SF6 | - 34.5sccm |
| - O2 | - 5.5sccm |
| - ICP forward power | - 500W |
| - RF forward power | - 5W |
| - He back pressure | - 10mT |
| - Pressure | - 0mT (2-6mT) |
| - Temperature | - -120 C |

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 08 June 2010 | Sam Bell |  |