Oxford ICP 2-step DRIE SOP

Scope

This document provides operating procedures for the Oxford 100 ICP 2-step DRIE.

Table of Contents

1.	Scope	2

2.	Table of Contents	2

3.	Reference Documents	2

3.1	Referenced within this Document	2

3.2	External Documents	2

4.	Equipment and/or Materials	3

5.	Safety	3

6.	Setup Procedures	3

6.1	Reserve and Enable tool in Coral	3

6.3	Vent Load Lock	5

6

6.4	Load Wafer	6

7.	Run Recipe	6

7.1	Load and Edit Recipe	6

7.2	Start Etch Process	8

8.	Shutdown Procedures	9

8.1	Vent Load Lock	9

8.2	Shutdown System	9

9.	Revision History	11

Figure 1, Liquid Nitrogen Tank	3

Figure 2, Chiller	3

Figure 3, Pump Control Page	4

Figure 4, Load Lock	4

Figure 5, Temperature Controller	5

Figure 6, Recipe Editor Page	7

Figure 7, Process Control Page	9

Figure 8, Etch Rate for 10 um Trenches	11

Figure 9, Etch Rate for 2 um Trenches	12

Figure 10, Etch Rate for 4 um Pillars	12

Figure 11, 2-step Etching SEM Analysis	13

Table 1, Deep Etch Recipe	6

Table 2, Smooth Etch Recipe	7

Reference Documents

Referenced within this Document

External Documents

None

Equipment and/or Materials

Wafer/Sample

Oxford 100

Liquid Nitrogen

Safety

Follow all Nanofab safety procedures.

Setup Procedures

Reserve and Enable tool in Coral

Etch-Oxford 100 DRIE (software Login is nanofabuser, pw nanofabuser)

Record all processing and characterization information in Coral

Set chuck temperature

Click on Process icon. See Figure 1, Pump Control Page.

Click on “Chamber 1” on the drop-down menu

Set the table temperature to -15 (lower left corner of screen) See Figure 1.

Press the pink start button, then immediately click stop (right next to it). See Figure 1.

- Figure 1. Set Chuck Temperature

Set point

Actual Temperature

Figure 2, Temperature Controller

Check the set point on the temperature controller to make sure the set point number reads -15 (see Figure 2).  The controller won’t function unless you initially set it way below your desired set point, and then change it to the actual set point later.

Find the unattached tube behind the main chamber.  Feel the output to make sure nitrogen is flowing out of it (see Fig 3).  This means the chuck is being cooled to the setpoint.

- Figure 3. N2 cooling flow tube

Now set the table temperature to your desired setpoint (typically 15C for most DRIE Bosch recipes) (lower left corner of screen) See Figure 1.

Press the pink start button, then immediately click stop (right next to it). See Figure 1.

Find the unattached tube behind the main chamber.  Feel the output to make sure nitrogen is still flowing out of it.

Vent Load Lock

Click on System icon. See Figure 4 Pump Control Page.

Click on “pumping” on the drop-down menu

Press stop button corresponding to the load lock mechanical pump (on the left side of the screen, do not touch the main chamber buttons on the right side of the screen). See Figure 4.

Press stop, then ok, then vent, then ok buttons. See Figure 1.

Wait for load lock to vent. The pressure will read above 600 Torr and you will hear hissing initially from the load lock lid.

System Button

Process Buttton

Load Lock Controls

Evacuate

Stop

Vent

Main Chamber Controls Do Not Touch

Figure 4, Pump Control Page

Figure 5, Load Lock

Press stop button. See Figure 4.

Load Wafer

Open load lock lid (See Fig 5)

Place wafer in load lock against the two pins on the transfer arm with the wafer facing the two pins. The wafer should touch the two pins. See Figure 5, Load Lock.

Close load lock lid.

Run Recipe

NOTE:	This recipe can be used to etch deep trenches or vias in 100 mm wafers. It is strongly advised that you do a test run on a practice wafer before working with your device wafer.

Load and Edit Recipe

Click on Process icon. See Figure 6, Pump Control Page.

Click on “Recipes” on the drop-down menu

Click on “New” and press ok to clear the current recipe

Click on “Load” and select the desired recipe (typically BB Bosch 12:5 15C 80:80 recipe for silicon DRIE), and click ok

The recipe steps will be displayed on the left side of the screen.

For Si DRIE set the number of etch cycles

Right click on “Repeat”

Click on “Repeat Step”

Enter the desired number of cycles to repeat the etch and deposition step for the Bosch process (typically ~0.67 um/cycle) but this varies greatly depending on the area of exposed silicon (can vary from ~0.4um/cycle to ~1.1 um/cycle).

Click OK

For other etch processes set the etch time

For other etch processes you can change the etch time by right clicking on the etch step and then changing the process time.

Do not change the gas flows or powers unless it is your own recipe and you know what you are doing.  Click Save when you are done editing to save the recipe.

Repeat Step Button Figure 6, Recipe Editor Page Figure 6, Recipe Editor Page Run Button Run Button

Figure 6, Load, Edit, and Run Recipe Page

The Bosch etch recipe is as follows:

| Table 1, Deep Etch Recipe |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Step | C4F8 | SF6 | Set Pressure | ICP Power | RF Power | Temp. (C) | Time |
| Flow C4F8 | 80sccm | 0 | 0mT | 0 | 0 | 15 | 5s |
| Strike Plasma | 80sccm | 0 | 0mT | 500W | 30W | 15 | 7s |
| Etch | 0 | 80sccm | 0mT | 500W | 25W | 15 | 12s |
| Deposit | 80sccm | 0 | 0mT | 500W | 10W | 15 | 5s |
| Repeat Etch & Deposit Steps |  |  |  |  |  |  |  |

In all steps the He backing is 10.0.

Start Etch Process

Make sure the chuck is at the correct temperature.  If it is too warm the Bosch process won’t work correctly.

Click the ‘Run’ button. See Figure 6.

NOTE:	The pumpdown and subsequent wafer transfer will begin automatically and the etch process will start when the wafer is loaded. The wafer will automatically transfer to the load lock when the process is complete.

Listen closely to the machine during the first steps.

If the machine makes a repetitive clicking sound, or the ICP power is greater than 10 or fluctuating wildly, you need to stop and restart the process.

Press Stop.

Return the wafer to load lock-YES

Click the Run button after you see that the wafer is inside the load lock to restart the process.

Monitor the parameters during the first few steps of the process.

Go to the Process Control page.

Monitor the Reflected Power. It should be less than 10. See Figure 7, Process Control Page.

Make sure the temperature cooling is working and stable +/- 1 degree C.

Reflected Power

Figure 7, Process Control PageReflected Power

Figure 7, Process Control Page

Shutdown Procedures

Vent Load Lock

Click OK when the software says “Process completed”.

Look inside load lock window to make sure wafer has transferred successfully from main chamber to load lock.

Vent load lock.

Press stop button corresponding to the load lock mechanical pump. See Figure 4.

Click ok when it says “Wafer has finished processing”.

Press vent button. See Figure 4.

Wait ~1 min. for the load lock to vent.

Pull up to open load lock lid. Do not force it open, once it is vented it should open easily.

Press stop button to stop venting.

Remove wafer.

Shutdown System

Set the chuck temperature to 30C.  This is important!  Otherwise the N2 dewar will run out overnight and you will be billed for a new tank.

Click on Process icon. See Figure 1, Pump Control Page.

Click on “Chamber 1” on the drop-down menu

Set the table temperature to 30 (lower left corner of screen) See Figure 1.

Press the pink start button, then immediately click stop (right next to it). See Figure 1.

Double check to make sure the bottom temperature setpoint on the controller reads 30.

- Figure 8. Temperature controller set to 30C

Run the Standby recipe to pump down the empty load lock.

Close load lock lid.

Load the “Standby” recipe

Run the Standby recipe

Important: Press ‘cancel’ when software asks for wafer name.  Do not click ok or enter a name or the system will get confused and think there is a wafer inside.

Enter run data, measurements, and disable tool in Coral.

- Figure 9. Run standby recipe to shutdown system

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 01 Jun 2010 | Sam Bell |  |
| 2 | 13 Dec 2021 | Brian Baker |  |