Oxford 80+ PECVD SOP

Scope

This document provides the operating procedure for plasma enhanced chemical vapor deposition with the Oxford 80+.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	1

3.1	Referenced within this Document	1

3.2	External Documents	2

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Setup Procedures	2

6.1	Record Information on Log Sheet	2

6.2	Turn on the Burn Box	2

6.3	Check Status of the Machine	2

6.4	Start Chamber Warm Up	3

7.	Deposition Procedures	4

7.1	Vent Chamber	4

7.2	Load Sample	4

7.3	Edit Recipe	4

7.4	Select Recipe	5

7.5	Start Deposition Process	5

7.6	Unload Sample	6

7.7	Pump Down Chamber	6

8.	Process Notes	7

8.1	Typical Film Characteristics	7

8.2	Process Summary	7

9.	Revision History	7

Figure 1, Oxford 80+ PECVD (Chamber A)	2

Figure 2, Oxford 80+ Chamber B (RIE)	3

Figure 3, Oxford Toolbar	4

Figure 4, Recipe Editor	5

Figure 5, Start Button	6

Table 1, Recipe Parameters and Results	7

Reference Documents

Referenced within this Document

Ecosys Burn Box SOP

External Documents

None

Equipment and/or Materials

Oxford 80+

Wafer/Sample

Tweezers

Safety

Follow all Nanofab safety procedures.

The stage can be very hot. Do not touch.

Setup Procedures

Record Information on Log Sheet

Record all requested setup and processing information on the log sheet.

Turn on the Burn Box

Follow the procedures in the Ecosys Burn Box SOP to turn on the burn box.

Figure 1, Oxford 80+ PECVD (Chamber A)

Open/Close Switch

Hoist Buttons

Check Status of the Machine

If the machine is not turned on, contact lab staff.

Ensure Oxford program is running on the PC.

If not, double-click the Oxford icon on the desktop.

If the switch is in Chamber A mode, skip to 6.4.

If the switch is in Chamber B mode. Do 6.3.4.1 - 6.3.4.3.

Abort pump down program if it is running.

Stop the turbo pump.

Open the Chamber B front panel.

Press the turbo pump stop button. See Figure 2, Oxford 80+ Chamber B (RIE).

Turn the chamber select switch to chamber A. See Figure 2.

Chamber Select Switch

Turbo Pump Stop Button

Figure 2, Oxford 80+ Chamber B (RIE)

Start Chamber Warm Up

Go to the recipe list.

Click the Select button. See Figure 3.

Select pump300c.rec recipe.

Click OK.

Press Start. See Figure 5.

Wait for the chamber to reach operating temperature. See Figure 5, Status Screen.

NOTE:	The status bar will read ‘Adjusting temperature’.

When the chamber has reached operating temperature, abort the pump down process.

Continue to section 7, Deposition Procedures.

Select Button

Display Button

Edit Button

Recipe List

Figure 3, Oxford Toolbar

Deposition Procedures

Vent Chamber

Select vent recipe.

Click Select. See Figure 3, Oxford Toolbar.

Select the vent recipe from the list.

Click OK.

Press start.

When the vent process is complete turn the selector knob to Open. See Figure 1.

Press and hold both Hoist buttons simultaneously. See Figure 1.

Swing the lid out of the way.

Load Sample

Use wafer tweezers to place sample on the stage.

Turn the selector knob to Close.

Press and hold both Hoist buttons to lower the lid.

You may have to use your forehead to align the lid as it closes.

Align lid with base.

Edit Recipe

Click the Edit button. See Figure 5.

Select Process from the drop down menu.

Go to File>Open.

Select a process recipe from the list.

Click on the process step to edit the parameters, as needed. See Figure 4.

Click OK.

Save the recipe.

Go to File>Save.

If you are using someone else’s recipe or making a new recipe, go to File>Save As to rename the recipe.

Figure 4, Recipe Editor

Select Recipe

Click Display button.

Choose Status from the drop down menu.

Click Select.

Choose your recipe from the list.

Click OK.

Start Deposition Process

Press Start Button.

NOTE:	At the end of most programs the machine will vent and be ready for unloading.

Monitor process parameters during deposition. See Figure 5, Status Screen.

Wait until status bar says Process complete.

Unload Sample

When the vent process is complete turn the selector knob to Open.

Press and hold both Hoist buttons.

Swing the lid out of the way.

Use wafer tweezers to pick up the sample.

Turn the selector knob to Close.

Press and hold both Hoist buttons to lower the lid.

You may have to use your forehead to align the lid as it closes.

Align lid with base.

Pump Down Chamber

Select pump down recipe.

Click Select.

Select pump21c.rec.

Click OK.

Press start.

Start Button

Figure 5, Status Screen

Process Notes

Typical Film Characteristics

| Table 1, Recipe Parameters and Results |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Recipe | Time | Pressure | Power | Thickness | Flatness |

Process Summary

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 10 June 2010 | Sam Bell |  |