Oxford 80+ RIE SOP

Scope

This document provides the operating procedure for the Oxford 80+ reactive ion etch.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	1

3.1	Referenced within this Document	1

3.2	External Documents	1

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Setup Procedures	2

6.1	Record Information on Log Sheet	2

6.2	Check Status of the Machine	2

7.	RIE Procedures	3

7.1	Vent Chamber	3

7.2	Load Sample	4

7.3	Edit Recipe	4

7.4	Select Recipe	5

7.5	Start Etch Process	5

7.6	Unload Sample	6

7.7	Pump Down Chamber	6

8.	Process Notes	6

8.1	Typical Etch Characteristics	6

8.2	Process Summary	6

9.	Revision History	6

Figure 1, Oxford 80+ RIE (Chamber B)	2

Figure 2, Turbo Pump Control Panel	3

Figure 3, Oxford Toolbar	3

Figure 4, Recipe Editor	4

Figure 5, Start Button	5

Table 1, Recipe Parameters and Results	6

Reference Documents

Referenced within this Document

None

External Documents

None

Equipment and/or Materials

Oxford 80+

Wafer/Sample

Tweezers

Safety

Follow all Nanofab safety procedures.

Setup Procedures

Record Information on Log Sheet

Record all requested setup and processing information on the log sheet.

Figure 1, Oxford 80+ RIE (Chamber B)

Chamber Select Switch

Hoist Buttons

Open/Close Switch

Check Status of the Machine

If the machine is not turned on, contact lab staff.

Ensure Oxford program is running on the PC. See Figure 3. If only the top bar is showing press display then status.

If not, double-click the Oxford icon on the desktop.

Check to make sure no other process is running. If another process is running contact lab staff.

If the switch is in Chamber B mode, skip to section 7 RIE Procedures.

If the switch is in Chamber A mode: Do 6.2.4.1 - 6.2.4.3.

Ensure that no program is running. Abort program if necessary.

Turn the chamber select switch to chamber B. See Figure 1.

Check that the light above turbo start button is lit. If it isn’t start the Turbo pump.

Start Button

Figure 2, Turbo Pump Control Panel

Open the Chamber B front panel.

Press the Turbo pump start button. See Figure 2.

Wait 5 minutes for turbo to spin up to full speed

RIE Procedures

Abort/Select/Start

Display

Edit

Figure 3, Oxford Toolbar

Vent Chamber

If machine is in pump down mode, click Abort to stop pump down.

Select vent recipe.

Click the Select button. See Figure 3.

Select the vent.rec recipe from the list.

Click OK.

Click Start. See Figure 5, Start Button.

- NOTE: The start button will not work unless the turbo is spinning at full speed.

Wait until Vac/gaspod interlock says fault.

DO NOT try to open the chamber lid unless it says fault.

NOTE:	The vent procedure takes 5 minutes.

Load Sample

When the vent process is complete turn the Chamber knob to Open. See Figure 1.

Press and hold both Hoist buttons simultaneously. See Figure 1.

Use wafer tweezers to place sample on the stage.

Turn the selector knob to Close.

Press and hold both Hoist buttons to lower the lid.

Figure 4, Recipe Editor

Edit Recipe

Click Edit. See Figure 3.

Choose Process from the drop down menu.

Go to File>Open.

Select a process recipe from the list.

Click on the process step to edit the parameters, as needed. See Figure 4, Recipe Editor.

Click OK

Save the recipe.

Go to File>Save.

If you are using someone else’s recipe or making a new recipe, go to File>New, suggested parameters are on sheep in front of binder

Select Recipe

Click Display button. See Figure 3, Oxford Toolbar.

Choose Status from the drop down menu.

Click Select.

Choose your recipe from the list.

Click OK.

Start Etch Process

Press Start Button.

Start Button

Figure 5, Start Button

NOTE:	At the end of most programs the machine will vent and be ready for unloading.

Observe CM gauge pressure, process gas flows, and RF power to assure etch program runs correctly.

Wait until Vac/gaspod interlock says fault and status bar says process complete.

If process compete but no “fault” reading, abort and run “vent” recipe.

Unload Sample

Turn the selector knob to Open when the vent process is complete.

Press and hold both Hoist buttons simultaneously.

Use wafer tweezers to pick up the sample.

Turn the selector knob to Close.

Press and hold both Hoist buttons to lower the lid.

Pump Down Chamber

Select pump down recipe.

Press Select.

Select pump down recipe.

Click OK.

Press Start.

Process Notes

Typical Etch Characteristics

| Table 1, Recipe Parameters and Results |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Recipe | Film Etched | Gas #1 Name | Gas #1 Flow | Gas #2 Name | Gas #2 Flow | Pressure | Power | Etch Rate |

Process Summary

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 10 June 2010 | Sam Bell |  |
| 2 | 24 Sep 2010 | Trevor Knowlton |  |