AW610 Rapid Thermal Process system SOP

Scope

This document contains information on how to operate the AccuThermo AW610.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	1

4.	Equipment Overview	2

5.	Safety	2

6.	Creating And Editing A Recipe	2

6.1	How to Create or Edit a Recipe	2

7.	Process A Wafer	3

7.1	How To Process a Wafer	3

8.	Daily Power-up & Power-down Procedures	3

9.	Warnings	4

10.	Revision History	4

Reference Documents

Referenced within this Document

AccuThermo AW610 Rapid Thermal Process System Operation Manual.

AccuThermo AW610 Rapid Thermal Process System Technical Manual.

External Documents

None

Equipment Overview

The AccuThermo RTP system is a rapid thermal processing system, which uses high intensity visible radiation to heat single wafers for short process periods of time at precisely controlled temperatures. The process periods are typically 1-600 seconds in duration, you must receive staff approval for a longer duration.

Safety

Follow all Nanofab safety procedures

Guidelines for Time-Temperature

The system is not capable of doing long anneals over extended periods of time, especially at higher temperatures. The following chart should serve as a broad guideline in determining allowable anneal time for the required temperature (roughly log dependence.) When in doubt, contact the staff before you run a new recipe.

Process A Wafer

How To Process a Wafer

Contact a staff member prior to making a recipe. Staff must create/approve ALL recipes

To log in, type a 5 in “User ID.” The username NanoFabUser should display automatically Type the password USR

Starting from the main menu, click “Process for Production”

On the left side of “DIR ID,” select USER

For “lot ID” select your username. If you do not have a username, type what you want your username to be and it will create it automatically

This is where all of your files will be saved

Select the recipe you plan to run

If more than one wafer is being processed, select the amount by clicking “wafer ID”

If your wafer is transparent or if you have wafer shards, proceed to section 7.1

Ensure the correct temperature measurement sensor is in place

Contact staff if the thermocouple needs to be installed or removed

Thermocouple is used for processes below 700°C

Pyrometer is used for processes between 550°C and 1150°C and the thermocouple must be removed above 700°C

Load your wafer by pulling the lever down to the unlock position and then pulling it forward

No wafer with photoresist, polymers or any other toxic or high vapor pressure material are allowed in the chamber

Once it is opened, you may carefully set your wafer on the quarts tray fingers

If using the thermocouple, insure it is in contact with the bottom of the wafer

Close the chamber. Confirm the lever is in the locked position

Turn lamp power on

Run the recipe by clicking on the “Start process” button

The computer will start the recipe automatically and display a graph on the monitor showing a live feed of data

The screen will show “PROCESS OVER” once the final step has ended. You can click “SAVE EXIT” or “STOP NOT SAVE”

Watch the LED display on the front of the RTP and wait until shows .050, which represents 50°C

This only works when using the thermocouple

Or wait 2 minutes minimum after the start of the cooldown and with metal tweezers, remove the wafer and place it onto foil for the final cooldown

Wait minimum of 4 minutes when using the susceptor

Open the chamber door and take out your sample

Close the chamber door and turn off the lamp power

Susceptor

How to use the Susceptor

If you are using a glass wafer, some other form of transparent wafer, or a part/shard of a wafer you must load it into the susceptor

If you set parts of a wafer on top of a regular wafer, it will move around during the process

Carefully open up the susceptor and set your wafer inside

Carefully close the susceptor with your wafer inside

When loading the susceptor into the chamber, ensure the numbers are facing upwards

Resume to step 6.1.6

Creating And Editing A Recipe

How to Create or Edit a Recipe

Contact a staff member prior to making a recipe. Staff must create/approve ALL recipes.

Starting from the main menu, click “Recipe” and select a pre-existing recipe

Click on the “Recipe Edit” button to display the recipe editor

Edit it (If a new recipe is being created, just change the recipe name field before saving it.)

If using the pyrometer as the temperature sensor device, then create an intensity step before the first ramp step. Set all the tuning parameters to 1.00 (gain, sensitivity and delay)

To avoid cracking the susceptor, the maximum ramp rate is 20°C/second

Click on “Recipe Head” to confirm RTP is set on the current calibrations

- Current calibration names:			Updated name, next cal

- PYRONAME	160930_3.WFR or SPT	______________________

- CHAMCALI	LAMPCAL.WFR or SPT	______________________

Save the recipe with .RCP or .SPT, SPT for susceptor recipes.

exit the recipe editor

All recipe’s must be approved by qualified staff before running a process

Daily Power-up & Power-down Procedures

Daily Power-Up

Ensure all gas valves are turned off

Verify there is no restriction to the process or cabinet exhaust and that they are set to the proper pressure and flow rate

Turn on the chamber cooling water

Turn on the quartz isolation tube cooling CDA/nitrogen.

Enable the lamps using the front panel Lamp Power switch. The green Lamp ON light indicates the lamps are enabled. You will also hear a "clunk" when the switch is turned on. This is the sound of the contactor engaging

Open the chamber door about 1 cm. This will ensure the quartz chamber will not over pressurize if there is a problem with the gas system

Turn on the gas valves

Daily Power-Down

Make sure it has been at least 5 minutes since the last use or until the chamber has cooled to room temperature

Turn off the gas valves

Turn off the oven unit LAMP POWER switch

Turn off the quartz isolation tube cooling CDA/nitrogen

- If any accessories are being used with the system, turn them off as necessary

Warnings

|  | WARNING - Long, high temperature processes may damage the AccuThermo RTP system - Make sure both water cooling and air cooling are on before running any process - Quartz is very delicate, handle with care |  |
| --- | --- | --- |

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 12 September 2016 | Tysen Heaton | Created SOP |