ProTemp Furnace SOP

Scope

This SOP provides instructions to operate the ProTemp Atmospheric Furnaces.

Table of Contents

1	Scope	1

2	Table of Contents	1

3	Reference Documents	2

3.1	External Documents	2

4	Equipment and/or Materials	2

5	Safety	2

6	General Information	2

7	Pre-Clean	4

7.6	Organic Clean Bath Preparation	4

7.7	SC-1 Bath Preparation	4

7.8	Oxide Removal Bath Preparation	5

7.9	SC-2 Bath Preparation	5

8	Furnace Operating Procedures	6

8.1	Log On	6

8.2	Enable Coral	6

8.3	Check Furnace Status	6

8.3.3	Review Current Status	6

8.4	Load Boat	8

8.4.1	Load Wafers	8

8.4.1.4	Solid Source Doping	8

8.4.1.5	Oxidation/Anneal	9

8.5	Initialize Graph	10

8.6	Change Oxidation Time	10

8.7	Download Recipe File	11

8.8	Select Recipe	12

8.9	Run Recipe	12

8.10	Wait for Run Completion	13

8.11	Unload Boat	13

8.12	Unload Wafers	13

8.13	Measure Monitor Wafers	13

8.14	Place System in Standby	13

8.15	Disable Tool in Coral	13

9	Solid Source Instructions	13

9.1	Install the desired boat	13

9.2	Dehydrate Sources	14

10	Process Notes	15

10.1	Process Summary	15

11	Revision History	16

Figure 1, Boat Fork	3

Figure 2, Solid Source Doping Loading Diagram	9

Figure 3, Oxidation/Anneal Loading Diagram	9

Table 1, Furnace Pre-Clean	5

Table 2, Log In Information	6

Table 3, Recipe File Location	11

Table 4, Oxidation Process Summary Data	15

Reference Documents

External Documents

SMBB Lab User Guide

Equipment and/or Materials

ProTemp Oxidation Furnaces

Quartzware

Boat fork

Monitor wafers

Phosphorous Doping:  <100> P-Type bare silicon

All Other Processes:  <100> N-Type or P-Type bare silicon

Filler wafers

Safety

Follow all Nanofab safety procedures.

Ensure all Cooldown cycles are complete before handling any quartzware, wafers, or samples.

|  | CAUTION |  |
| --- | --- | --- |
|  | All recipes include the necessary cool down time to ensure the quartzware, wafers, and samples are cool enough to be touched and handled. Do not touch or handle quartzware, wafers, or samples unless the current run is COMPLETE. |  |
|  | CAUTION |  |

General Information

Ensure all samples are authorized to be loaded into a furnace.

|   | !!! WARNING !!! |   |
| --- | --- | --- |
|    | No photoresist, kapton tape, metals, glass, or other foreign material is allowed in any furnace. Only quartz, silicon, silicon carbide, silicon dioxide, silicon nitride, and/or Polysilicon are allowed without PRIOR approval of lab staff. |    |
|   | !!! WARNING !!! |   |

Only the following substrates or samples are allowed in the Clean Ox furnace:

- Bare silicon wafers that have received NO prior processing

- Wafers with an oxide that was thermally grown in the Clean Ox furnace

- Wafers with LPCVD Undoped Polysilicon deposition

- Wafers with LPCVD Nitride deposition

The following substrates or samples are allowed in the Doped Ox furnace:

- All samples approved for processing in the Clean Ox furnace

- Wafers that have been processed through PECVD Oxide deposition

- Wafers that have been processed through PECVD Polysilicon deposition

- Wafers that have been processed through PECVD Nitride deposition

- Wafers that have been processed through LPCVD Doped Poly Deposition

- Wafers that have been processed through LPCVD LTO Deposition

- Wafers that have been processed through LPCVD PSG Deposition

- Wafers that have been doped by any method

For any substrates or samples not listed above, contact Staff to determine the authorized furnace.

Ensure all samples are clean, dry, and particle-free.

Place quartzware on a quartz plate, silicon carbide cantilever paddle, or a cleanroom wiper.

Do NOT directly place quartzware on stainless steel.

Do NOT touch boats or boat covers with your hands.

Use the boat fork (see Figure 1, Boat Fork) to move boats and boat covers.

Insert the tines of the fork into the pick-up tubing on the boat or boat cover.

Figure 1, Boat Fork

Before touching a quartz baffle, put on a second pair of clean blue nitrile gloves.

Solid Source wafers that have been out of the furnace more than one (1) hour need a dehydration bake per section 9 Solid Source Instructions.

Pre-Clean

For new, unused bare silicon wafers, a Pre-Diffusion Clean prior to processing through any furnace cycle is recommended.

For all other samples/substrates/wafers, a Pre-Diffusion Clean prior to processing through any furnace cycle is REQUIRED.

Complete the Pre-Diffusion Clean in the Pre-Diffusion wet bench.

The dedicated baths should be used for whole wafers, since the SC-1 bath includes a megasonic clean for improved particle removal.  However, for partial wafers the specified mixtures may be used in smaller beakers/containers, if necessary.

Refer to Table 1, Furnace Pre-Clean for the correct mixtures, temperatures, and times.

If the Organic Clean bath temperature is above 100°C and has not been used within the last 2 hours, add 100ml of Hydrogen Peroxide.

Wait at least 5 minutes before using the bath.

Organic Clean Bath Preparation

If using an alternate bath container, ensure the container is compatible with temperatures up to 200°C (e.g., pyrex, quartz, teflon).

Add the necessary amount of Sulfuric Acid to an empty container.

Turn on the heater.

Wait for the heater to exceed 100°C.

Add the necessary amount of Hydrogen Peroxide.

Wait until the bath temperature has reached setpoint.

Wait at least 5 minutes before using the bath.

SC-1 Bath Preparation

If using an alternate bath container, ensure the container is compatible with temperatures up to 100°C (e.g., pyrex, quartz, teflon).

Add the necessary amount of Ultra Pure Water (UPW) to an empty container.

Add the necessary amount of Ammonium Hydroxide.

Turn on the heater.

Wait for the heater to exceed 65°C.

Add the necessary amount of Hydrogen Peroxide.

Wait until the bath temperature has reached setpoint.

Wait at least 5 minutes before using the bath.

Oxide Removal Bath Preparation

If using an alternate bath container, ensure the container is compatible with Hydrofluoric Acid (e.g., plastic, teflon).

Add the necessary amount of Ultra Pure Water (UPW) to an empty container.

Add the necessary amount of Hydrofluoric Acid.

Wait at least 5 minutes before using the bath.

SC-2 Bath Preparation

If using an alternate bath container, ensure the container is compatible with temperatures up to 100°C (e.g., pyrex, quartz, teflon).

Add the necessary amount of Ultra Pure Water (UPW) to an empty container.

Add the necessary amount of Hydrochloric Acid.

Turn on the heater.

Wait for the heater to exceed 65°C.

Add the necessary amount of Hydrogen Peroxide.

Wait until the bath temperature has reached setpoint.

Wait at least 5 minutes before using the bath.

| Table 1, Furnace Pre-Clean |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Step | Description | Container Type | Ratio1 | Chemical | Quantity (ml) | Temp (°C) | Time (mm:ss) |
| 1 | Organic clean (pirahna) | Pyrex Quartz Teflon | 5 1 | Sulfuric Acid Hydrogen Peroxide, 30% | 2500 500 | 120 | 10:00 |
| 2 | Rinse |  |  | Ultra Pure Water (UPW) |  |  | 5:00 |
| 3 | SC-1 | Pyrex Quartz Teflon | 5 1 0.4 | Ultra Pure Water (UPW) Ammonium Hydroxide Hydrogen Peroxide | 6250 1250 500 | 75 | 10:00 |
| 4 | Rinse |  |  | Ultra Pure Water (UPW) |  |  | 5:00 |
| 5 | Oxide Removal | Plastic Teflon | 50 1 | Ultra Pure Water (UPW) Hydrofluoric Acid, 49% | 2000 40 |  | 01:00 |
| 6 | Rinse |  |  | Ultra Pure Water (UPW) |  |  | 5:00 |
| 7 | SC-2 | Pyrex Quartz Teflon | 6 1 1 | Ultra Pure Water (UPW) Hydrochloric Acid, 37% Hydrogen Peroxide | 3900 650 650 | 75 | 10:00 |
| 8 | Rinse |  |  | Ultra Pure Water (UPW) |  |  | 5:00 |
| 9 | Spin-Rinse Dry |  | As programmed |  |  |  |  |
| 1Maintain this ratio when using the alternate bath containers. 2If the Organic clean bath temperature is greater than 100°C and has been unused for more than 2 hours, add 100ml of Hydrogen Peroxide prior to use. |  |  |  |  |  |  |  |

Furnace Operating Procedures

Log On

From the ProTemp host computer, select the Tymplex program.

If necessary, open the Tymplex program.

From the  0 Overview  screen, click the  Log In  button.

Type the Username and password (in CAPS) for the applicable furnace.  (See Table 2, Log In Information.)

From the  User Login …  dialog box, click the  Log In  button.

Overview Tab

Log In Button

| Table 2, Log In Information |  |  |
| --- | --- | --- |
| Furnace | Clean Ox | Doped Ox |
| Username | CLEANOX | DOPEDOX |
| Password | CLEANOX | DOPEDOX |

Enable Coral

Enable the furnace in Coral.

Check Furnace Status

Select the main tab for the desired furnace.

Select the Operate subtab.

Review Current Status

If the current recipe is “IDLE WITH COMPLETED SAMPLES ON BOAT”, samples from the previous run have not been removed from the boat.

If the current recipe is “IDLE WITH COMPLETED SAMPLES ON BOAT”, do the following:

If there are no instructions from the previous user, contact that user before removing any samples and monitor wafers.

If the instructions from the previous user are “Do Not Disturb”, do not proceed.

If the instructions from the previous user authorize you to handle their samples, continue.

Operate subtab

Doped Ox Furnace main tab

Blue dot

Current Recipe

If there is a blue dot next to the  Step  button, the furnace is on Hold.

Click the  Run  button.

Wait for the current run to complete.

The  Alarm Silence  button will be replaced by a yellow  *** Complete ***  button with a yellow dot.

Yellow Dot

Reset button

If there is a yellow dot next to the  *** Complete ***  button, the current recipe has completed.

Click the  *** Complete ***  button.

Click the  Reset  button.

If the cantilever is in and there is a blue dot next to the  Reset  button, download and run the “BOAT OUT” recipe per paragraphs 8.7 Download Recipe File through 8.9 Run Recipe.

If the cantilever is in and there is a yellow dot next to the  *** Complete ***  button, click the  *** Complete ***  button.

Download and run the “BOAT OUT” recipe per paragraphs 8.7 Download Recipe File through 8.9 Run Recipe.

If the cantilever is out, the furnace is ready for use.

If samples from another member are loaded on the boat AND he/she has authorized you to handle their samples, follow their instructions to remove the samples.

Load Boat

Load Wafers

|   | !!! WARNING !!! |   |
| --- | --- | --- |
|    | No photoresist, kapton tape, metals, glass, or other foreign material is allowed in the furnace. Only quartz, silicon, silicon carbide, silicon dioxide, silicon nitride, and/or Polysilicon are allowed without PRIOR approval of lab staff. |    |
|   | !!! WARNING !!! |   |

Unless otherwise noted, load all wafers with the polished side toward the door of the furnace.

Load the wafers with the wafer flat up.

Load one wafer per slot.

Solid Source Doping

Solid Source wafers that have been out of the furnace more than one (1) hour need a dehydration bake per section 9 Solid Source Instructions.

The solid source wafers should always have a filler or sample wafer in the adjacent slots – even when stored.

Remove the necessary filler wafers (adjacent to the solid source wafers) and place them in empty slots at either end of the boat.

Load all sample and monitor wafers with the polished side (side to be doped) toward the solid source wafer.

Load a monitor wafer onto the boat in slot 14 (see Figure 2, Solid Source Doping Loading Diagram).

Monitor wafers for Phosphorous Doping are <100> P-Type bare silicon.  Monitors wafers for Boron Doping are <100> N-Type bare silicon.

Load sample wafers in slots 9-13 and/or 15-18, starting at the center.

DOOR

E = Empty/Extra Slot	F = Filler wafer	M = Monitor wafer	S = Sample Wafer

| Quartz Baffle |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Quartz Baffle |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Slot # Wafer |  | 1 E |  | 2 E |  | 3 E |  | 4 F |  | 5 F |  | 6 F |  | 7 F |  | 8 E |  | 9 S |  | 10 S |  | 11 S |  | 12 S |  | 13 S |  | 14 M |  | 15 S |  | 16 S |  | 17 S |  | 18 S |  | 19 E |  | 20 F |  | 21 F |  | 22 F |  | 23 F |  | 24 E |  | 25 E |  | 26 E |  |  | Slot # Wafer |

Figure 2, Solid Source Doping Loading Diagram

Load filler wafers in slots 4-7, 20-23, and all unused sample slots.

Oxidation/Anneal

Load a monitor wafer onto the boat in slot 14 (see Figure 3, Oxidation/Anneal Loading Diagram).

Monitor wafers are bare silicon.  Samples that are bare silicon may be used as monitors.

Load sample wafers in slots 8-13 and/or 15-20, starting at the center.

Load filler wafers in slots 1-7, 21-25, and all unused slots.

DOOR

F = Filler wafer	M = Monitor wafer	S = Sample Wafer

| Quartz Baffle |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Quartz Baffle |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Slot # Wafer |  | 1 F |  | 2 F |  | 3 F |  | 4 F |  | 5 F |  | 6 F |  | 7 F |  | 8 S |  | 9 S |  | 10 S |  | 11 S |  | 12 S |  | 13 S |  | 14 M |  | 15 S |  | 16 S |  | 17 S |  | 18 S |  | 19 S |  | 20 S |  | 21 F |  | 22 F |  | 23 F |  | 24 F |  | 25 F |  |  | Slot # Wafer |

Figure 3, Oxidation/Anneal Loading Diagram

Polished Side

Ensure there is a sample, monitor, or filler wafer in every slot.

Initialize Graph

Select the  Setup  subtab.

Ensure the  Enable Datalog  check box is checked.

Press the  Create New Datalog  File button.

Setup subtab

Enable Datalog checkbox

Create New Datalog File button

At the  Creating new datalog file …  box, click  OK .

Change Oxidation Time

From the desktop, open the “Change OxTime” utility.

Enter/select your Coral ID.

Following the prompts, change the oxidation time, if needed.

Download Recipe File

To download a new recipe file, select the  Download  subtab.

Click the  Download File to Tymkon  button.

From the dialog box, locate and select the correct recipe file (see Table 3, Recipe File Location).

From the  Ready to Download  dialog box, click the  Begin Download  button.

Download File to Tymkon button

Download subtab

This text changes from red to black when download is complete.

Wait for the recipe to complete downloading.  (Sections will change from red to black text.)

| Table 3, Recipe File Location |  |  |  |
| --- | --- | --- | --- |
| Furnace | Path | Recipe Type | Recipe File |
| CleanOx | [MainPath]\CleanOx\CleanOx Recipes | Dry Oxidation Wet Oxidation | Clean DryOx Recipes.mdb Clean WetOx Recipes.mdb |
| DopedOx | [MainPath]\DopedOx\DopedOx Recipes | Anneal Boron Doping Dry Oxidation Phosphorous Doping Wet Oxidation | Doped Anneal Recipes.mdb Boron Solid Source Recipes.mdb Doped DryOx Recipes.mdb Phos Solid Source Recipes.mdb Doped WetOx Recipes.mdb |
| [MainPath] = C:\Users\Public\Documents\TymkonTools\UserData |  |  |  |

Select Recipe

Select the  Operate  subtab.

Click the  Select Recipe  pull down box.

Select the desired recipe.

Select Recipe pull down box

Select Recipe button

Activate the recipe by clicking the  Select Recipe # (recipe name)  button.

Run Recipe

From the  Operate  subtab, click the  Run  button.

Run button

Wait for Run Completion

Wait for the cooldown cycle to complete.

The boat will be out, the current recipe will be “REMOVE SAMPLES” and the current cycle will be “REMOVE SAMPLES”.  If this recipe is not reset within a programmed amount of time, the system will start the “IDLE WITH COMPLETED SAMPLES ON BOAT” and the load will push back into the furnace.

Click the  Reset  button.

At the  Reset the current recipe ?  dialog box, click the  OK  button.

Unload Boat

Ensure the run has completed.

|  | CAUTION |  |
| --- | --- | --- |
|  | All recipes include the necessary cool down time to ensure the quartzware, wafers, and samples are cool enough to be touched and handled. Do not touch or handle quartzware, wafers, or samples unless the current run is COMPLETE. |  |
|  | CAUTION |  |

Unload Wafers

Remove the monitor wafers and samples from the boat.

Measure Monitor Wafers

Measure the thickness at 5 sites on each monitor wafer.

Place System in Standby

Select the “VACUUM IDLE” recipe.

Run the recipe.

The cantilever will go in, the system will pump down, and go on hold.

Disable Tool in Coral

Disable the furnace in Coral.

Record the thickness measurements.

Solid Source Instructions

Install the desired boat

Ensure the unwanted quartz boat on the cantilever is sufficiently cool.

The programmed cooldown cycle will be finished and the host computer will show  *** Complete *** .

Using the boat fork, carefully remove the unwanted boat and place it on a quartz plate for storage.

Using the boat fork, carefully place the desired boat on the cantilever.

Dehydrate Sources

If the solid sources have been out of the furnace for more than 1 hour, dehydrate the sources.

Using the boat fork, place the boat with the solid sources on the cantilever.

Ensure each slot adjacent to a solid source wafer has a filler wafer.

Process the “DEHYDRATE SOURCES” recipe.

When completed, the sources are ready to be used for doping.

Process Notes

Process Summary

Table 4, Oxidation Process Summary Data lists the basic information for the various processes.  Contact staff for help selecting a recipe and oxidation time.

| Table 4, Oxidation Process Summary Data |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Furnace | Recipe | N2 (slm) | O2 (slm) | N2/TLC (slm) | Temp (°C) | Approx. Time (min) | Thick (Å) |
| CleanOx, Doped Ox | DRY OX (All temps) |  | 4 |  | 1000 | 60 | 550 |
|  |  |  |  |  | 1050 | 27 31 53 78 | 400 500 750 1000 |
|  | CLEAN OX (All temps) |  | 3.5 |  | 850 | 100 | 800 |
|  |  |  |  |  | 950 | 32 57 80 | 1000 1500 2000 |
|  |  |  |  |  | 1000 | 30 41 55 | 1500 2000 2500 |
|  |  |  |  |  | 1050 | 30 48 102 300 600 | 2200 3000 5000 10000 14000 |
| Doped Ox | ANNEAL (All temps) | 5 |  |  |  |  |  |
|  | BORON DOPING | 8 |  |  |  |  |  |
|  | PHOS DOPING | 8 |  |  |  |  |  |
| Note: Oxidation times can be approximated by this formula: New Time=Known Time*New Thickness2(Known Thickness)2 |  |  |  |  |  |  |  |

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 3 | 10 Jun 19 | T. Olsen | Update document to new format. |
| 2 | 29 Jun 18 | T. Olsen | General update and re-order sub-sections. |
| 1 | 28 Mar 16 | T. Olsen | Initial Release |