Suss Aligner SOP

Scope

This document provides the procedure for operating the Suss Aligner.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	2

3.1	Referenced within this Document	2

3.2	External Documents	2

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Setup Procedures	2

6.1	Turn on UV Lamp Supply	2

6.2	Turn on Machine	3

6.3	Insert Mask	3

6.4	Set Exposure and Gap	3

7.	Alignment Procedure	4

7.1	Load Wafer	4

7.2	Align Wafer	4

7.3	Expose	5

8.	Shutdown Procedure	5

9.	Process Notes	6

9.1	Test Run Results	6

10.	Revision History	6

Figure 1, UV Power Supply	3

Figure 2, Left Side Controls	3

Figure 3, Right Side Controls	4

Table 1, Test Results	6

Reference Documents

Referenced within this Document

None

External Documents

None

Equipment and/or Materials

Suss Aligner

Mask

Wafer/Sample

Safety

Follow all Nanofab safety procedures.

Setup Procedures

Turn on UV Lamp Supply

Press Start. See Figure 1, UV Power Supply.

NOTE:	Display flashes “cold” until the lamp is warmed-up and ready

Start Button

Figure 1, UV Power Supply

Make sure LED is on CI1 (Constant Intensity is set to 10mW/cm^2)

Power Switch

Figure 2, Left Side Controls

Turn on Machine

Turn on black knob with 0/1 written on it, on left side of machine. See Figure 2, Left Side Controls.

Insert Mask

Lift up mask holder.

Insert mask into mask holder (mask goes between pins, chrome side up).

Put mask holder down and push gently to snap into place.

Wait for system to boot up.

Turn on Mask Vacuum (knob on left-hand side).

Press enter for regular alignment.

Set Exposure and Gap

Press edit twice. See Figure 3, Right Side Controls.

Arrow to Exp-Type

Press edit and use arrows to set desired type.

Soft for features greater than 5 microns

Hard for features less than 5 microns

Vacuum for features less than 2 microns (wafers only)

Prox-1 and Prox-2 for features greater than 15 microns

Press enter. See Figure 3.

Arrow to Algn-Sep for the gap setting

Press edit, arrow to desired value, then press enter

Arrow to On Time

Press edit

Use arrows to adjust time, then press enter

Press enter again to return to main menu

Select Save and press enter

NOTE:	Your program should appear on screen

Alignment Procedure

Load Wafer

Pull slider out (when it is all the way out, the vacuum is off)

Insert water

Make sure the cross hairs are lined-up on the front of the machine (glass slides protruding from the front center of the machine, just below the slider).

Rotate knobs to align

Push slider in

Press ‘Load Mask’ (light should go off). See Figure 3.

Press ‘First Exp’ (so that the light goes off). See Figure 3.

Press ‘Align Check’ (so that the light goes on). See Figure 3.

Press ‘Start’ (microscope moves over). See Figure 3.

Figure 3, Right Side Controls

Align Wafer

To align microscope to the mask:

Align Check can be on or off, Live must be on, TSA on left side must be on

Use arrow buttons on the left side to move the objectives

The outside knobs on the microscope are for distance alignment

The small inner knobs on the microscope move the objectives up/down independently

The knob coming out at an angle from the microscope is for rotation (theta)

To align wafer to the mask:

Use the large knobs on the middle of the machine. Rotate knobs to align. (right y, leftx, back left=theta)

Course Focus is the back vertical knob on top. RESET is used for ESC.

Press ‘Cont/Sep’ (light should go off)

Expose

Press ‘Exp’ twice (microscope moves off and exposure lamp moves over). See Figure 3.

Slider pops out

NOTE:	If vacuum is too low, an error can appear at the end of the exposure, reading briefly “no vacuum, pull slider”, followed by “z-axis error”. Pressing ‘RESET’ will allow for unloading the wafer.

Shutdown Procedure

Remove wafer and push slider in.

Turn off Mask Vacuum.

Press Load Mask (light goes on)

Unload Mask

Turn off machine

Leave UV light on.

Process Notes

Test Run Results

| Table 1, Test Results |  |  |  |  |
| --- | --- | --- | --- | --- |
| Type of Contact | Exp. Time (s) | Aln-Sep (microns) | Development Time (s) | Results |
| Soft | 5 | 15 | 45 | poor; width of border of micro '5' feature= 9.10 microns |
| Soft | 7 | 15 | 45 | poor |
| Hard | 6 | 15 | 45 | good |
| Hard | 5 | 15 | 45 | good; width of border= 2.35 microns |
| Hard | 4 | 15 | 45 | good; width of border= 2.82 microns |
| Vacuum | 5 | 15 | 45 | worse than Hard, better than Soft; width of border= 5.49 microns |
| Vacuum | 5 | 15 | 60 | better than the previous vacuum test-run; width of border= 2.98 microns |

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 12 Jan 2010 | Sam Bell |  |