## Sheet1
Denton 635 Baseline
Chamber integrity | value | [7 empty] | date
Baseline pressure, Torr | 4e-08 | [7 empty] | 2016-12-15
Chamber Pumpdown , Atm to base | tbd
CM zero at base pressure | 0.6 | [7 empty] | 2016-12-15
Compare to "golden" CM's at 10mT, 20 mT | tbd
Leak up from base (2 min) Torr | 1.2e-06 | [7 empty] | 2016-12-15
| Pre | Post | rate
leak rate (10 min) @ 5 mT | 5.3 | 5.3 | 0 | [5 empty] | 2016-12-15
LL Pumpdown (time to 0.1T) | 2:02 | Autopump | 2:15 | [5 empty] | 2016-12-15
LL leakrate @ 0.01 T | tbd
Cryo low temp K, (LED) | 12.5 | [7 empty] | 2016-12-15
Cryo warm temp K, (LED), off and purged | 304 | [7 empty] | ~ 12/7/16
MFC's | Flow rating | Cal factor | Chamber pressure @ 10% CM gauge hi-vac gauge |  | Chamber pressure @ 50% CM gauge hi-vac gauge |  | Chamber pressure @ 90% CM gauge hi-vac gauge
1- Ar, SCCM 10, 50, 90 | 100 | 1.39 | 0.8 | 2.3 e-4 | 1.7 | 1.0 e-3 | 2.3 | 1.6 e-3
2- no gas | 100 | 1.39
3- N2, SCCM 5, 25, 45 | 51 | 1 | 0.7 | 9.0 e-5 | 1.2 | 4.0 e-4 | 1.5 | 7.0 e-4
4- O2, SCCM 5, 25, 45 | 51 | 0.99 | 0.7 | 7.8 e-5 | 1.2 | 4.0 e-4 | 1.6 | 7.0 e-4
Throttle Valve | POS 200 | POS 300 | POS 400 | POS 500 | POS 600 | POS 700 | POS 800 | POS 900
Ar 90 sccm - Chamber pressure mT | 54.9 | 39.8 | 9.8 | 4.1 | 2.9 | 2.6 | 2.4 | 2.3 | 2016-12-15
RGA
635 System gauge base pressure, Torr | 4.0 e-8 | [7 empty] | 2016-12-15
Mass 18, H2O | 1.3 e-7 | [7 empty] | 2016-12-15
Mass 28, CO or N2 | 5.4 e-9 | [7 empty] | 2016-12-15
Mass 40, Ar | 1.7 e-10 | [7 empty] | 2016-12-15
Mass 44, CO2 | 3.0 e-9 | [7 empty] | 2016-12-15
Mass 64, PR | n/r | [7 empty] | 2016-12-15
| Cath 1 | Cath 2 | Cath 3 | Cath 4
Targets | Au_4 | Cr_4 | Al_3 | Ti_3 |  |  |  |  | 2016-12-15