Nanospec 3000 SOP

Scope

This document provides procedures for measuring film thickness with the Nanospec 3000.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	2

3.1	Referenced within this Document	2

3.2	External Documents	2

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Setup Procedures	2

6.1	Load Sample	2

6.2	Select Program	2

6.3	Calibration	3

7.	Measurement Procedures	4

7.1	Film Thickness Measurement	4

8.	Revision History	5

Figure 1, Measurement Screen	3

Figure 2, Microscope Components	4

Figure 3, Programs List	5

Reference Documents

Referenced within this Document

None

External Documents

None

Equipment and/or Materials

Nanospec

Wafer/Sample

Calibration Wafer

Safety

Follow all Nanofab safety procedures.

Setup Procedures

Load Sample

Slide out microscope stage.

Place sample in the center of the stage, polished side up.

Slide in the stage.

Focus on the sample. See Figure 2.

Select Program

If in the measurement screen, check the top to see which program is selected.

If it is the program you need;

Press New Test. See Figure 1, Measurement Screen.

Enter your sample number.

Skip to section 7

If you need a different program proceed to step 6.2.2

Click Calib. See Figure 1.

Select a program from the list. See Figure 3, Programs List.

When prompted by the yellow filter warning, turn the yellow filter on or off. See Figure 2.

If the warning says “Yellow filter”, turn on the yellow filter.

If the warning says “No filter”, turn off the yellow filter.

Calibration

If prompted to “Measure New Reference?” , click Yes.

When prompted by the dark reference measurement,

Move the sample out of the way of the microscope reticule.

Click Ok.

When prompted to focus on reference sample,

Place calibration wafer on top of your wafer, polished side up.

Move calibration wafer under the microscope reticule.

Focus on the calibration wafer.

Click Ok.

Figure 1, Measurement Screen

Yellow Filter Switch

Focus Wheel

X/Y Knob

Figure 2, Microscope Components

Measurement Procedures

Film Thickness Measurement

Use X and Y knob to move sample to measurement point. See Figure 2.

Focus on the sample.

Click the Meas button to take a measurement. See Figure 1, Measurement Screen.

Repeat as necessary.

Figure 3, Programs List

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 17 June 2010 | Sam Bell |  |