TM Vacuum Sputter System SOP



1. Scope
   1. This document provides the procedures and requirements to operate the TM Vacuum Sputter system.

1. Table of Contents

[1 Scope 1](#_Toc290639668)

[2 Table of Contents 1](#_Toc290639669)

[3 Reference Documents 2](#_Toc290639670)

[3.1 Referenced within this Document 2](#_Toc290639671)

[3.2 External Documents 2](#_Toc290639672)

[4 Equipment and/or Materials 2](#_Toc290639673)

[5 Safety 2](#_Toc290639674)

[6 Setup Procedures 2](#_Toc290639675)

[7 Load Samples 2](#_Toc290639676)

[8 Pump Down the Load-lock Chamber 6](#_Toc290639677)

[9 Transfer Carousel to the Main Chamber 7](#_Toc290639678)

[10 DC Sputter 8](#_Toc290639679)

[11 Transfer of Carousel back to Load-Lock 14](#_Toc290639680)

[12 Unload Samples from the Load-lock 14](#_Toc290639681)

[13 Troubleshooting 15](#_Toc290639682)

[13.1 Plasma does not Ignite 15](#_Toc290639683)

[14 Process Notes 16](#_Toc290639684)

[14.1 TMV Parameters 16](#_Toc290639685)

[14.2 TMV Target Positions 16](#_Toc290639686)

[15 Revision History 17](#_Toc290639687)

[Figure 1, TMV Chambers 3](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639644)

[Figure 2, Ion Gauges 3](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639645)

[Figure 3, Control Mode 3](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639646)

[Figure 4, Load Lock Control Panel 4](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639647)

[Figure 5, Sample Loading 4](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639648)

[Figure 6, Chuck Positions 5](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639649)

[Figure 7, Handle (Secured Position) 5](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639650)

[Figure 9, Load Lock Vacuum Panel 6](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639652)

[Figure 11, Main Menu 7](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639654)

[Figure 12, Gate Valve Buttons 7](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639655)

[Figure 13, Automatic Control Panel 8](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639656)

[Figure 14, Substrate Index Rotation Control Window 9](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639657)

[Figure 15, MKS Pressure Controller 10](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639658)

[Figure 16, Sample Rotation Selector 10](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639659)

[Figure 17, Power Supply/Cathode Selection Screen 11](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639660)

[Figure 18, DCGs 11](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639661)

[Figure 19, Plasma Supply Control 12](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639662)

[Figure 20, Cathode Shutter Control 13](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639663)

[Figure 21, TMV Target Positions 16](file:///C:\Documents%20and%20Settings\Labstaff\Desktop\nanodocs\Microfab-Deposition\TM-Vacuum%20Sputtering%20System\TMV%20SOP,%20100507%20rev1.docx#_Toc290639664)

[Table 1, Power Setpoints 12](#_Toc290639665)

[Table 2, Cathode Materials 13](#_Toc290639666)

[Table 3, Sputter Parameters 16](#_Toc290639667)

1. Reference Documents

Referenced within this Document

* + 1. None

External Documents

* + 1. None

1. Equipment and/or Materials
   1. TMV Sputtering System
   2. Wafer
   3. Kapton Tape
   4. Nylon Screws
   5. Flashlight
   6. Crowbar
2. Safety
   1. Follow all Nanofab safety procedures.
   2. Place carousel only on clean stainless steel surfaces.
   3. Ask for assistance moving carousel if too heavy to remove/mount alone.
   4. Do NOT use a chair/stool with wheels to stand on while attempting to remove/mount the carousel.
   5. Use only kapton tape or plastic screws to secure samples to the rotating chuck.
3. Setup Procedures
   1. Reserve and enable TMV system in Coral.
4. System will not vent without being enabled.
   1. Record all information on log sheet.
5. Load Samples
   1. Use flash light to check that the carousel is in the load-lock chamber. See *Figure 1, TMV Chambers*.
   2. Check to see that the following buttons are lit on the control panel.



Load Lock

Main Chamber

Figure , TMV Chambers

[A] load-lock cryo-pump (yellow button)

[B] load-lock high-vac valve (green button)

[C] main chamber cryo-pump (yellow button)

[D] main chamber high-vac valve (green button)

* + 1. **If not lit**, DO NOT turn on the ion gauge for either the main chamber or the load-lock, and contact the fab staff.
  1. **If there is a reading on top left IG display**, press Chamber Gauge button to turn it OFF. See *Figure 2, Ion Gauges*.
  2. Check reading A on Main Chamber ion gauge; the reading should be 0.0. See *Figure 2*.



Chamber IG

Load Lock IG

Gauge Buttons

Figure , Ion Gauges

Reading A

* 1. **If load-lock Ion Gauge is on (top right IG display),** press Gauge to turn itOFF**.** See *Figure 2*.
  2. Turn CONTROL MODE key to AUTO position. See *Figure 3, Control Mode*.

1. The AUTO light should be on.
   1. If load lock AUTO PUMP is on, turn it OFF. See *Figure 4, Load Lock Control Panel*.





Figure , Control Mode

* 1. Press AUTO VENT on the Load-lock panel. See *Figure 4*.
     1. If AUTO VENT does not work, (light does not come on, no gas sound), press TRANSFER GATE VALVE CLOSE, then try AUTO VENT again.
  2. If AUTO VENT still does not work, put the key in MANUAL mode.
     1. Close the load-lock high-vac valve (light off).
     2. Open the load-lock vent valve (light on).



Load Lock AUTO PUMP Button

Load Lock AUTO VENT Button

Figure , Load Lock Control Panel

* + 1. Close the load-lock vent valve after the door opens (light off).
    2. Return the key to AUTO mode.
  1. Once the load lock door pops open, press AUTO VENT again to stop the nitrogen flow.

1. Pay attention to glove cleanliness when handling the carousel so as not to contaminate your sample.
2. If you are only using 1 or 2 chucks, you may load the wafers onto the chuck(s) without removing the carousel.
   1. If you are using more than 2 chucks, remove the carousel from the load-lock.
      1. Lift the carousel off the pins.
      2. Rotate the carousel 1/16 of a turn.



Carousel

Nylon Screw

Chuck

Wafer

Figure , Sample Loading

* + 1. Slide the carousel down off of it’s mount.
  1. Load samples on rotating chucks. See *Figure 5*.

1. Only 2 of the chucks are able to rotate. If you need chuck rotation, load your samples on the chucks with small gears on the back.
   * 1. Attach samples on the chuck with the nylon screws or kapton tape.

1

7

8

6

5

4

3

2

Sample Holders

Front of TMV

Chuck Positions as Loaded into the Load Lock

Figure , Chuck Positions

1. The nylon screws are located in the small black toolbox.
   1. Note the position of your samples on the carousel. See *Figure 6, Chuck Positions*.
   2. Mount the carousel in the load lock (ask for help from others).




Figure , Handle (Secured Position)

* + 1. Slide the carousel up over the mounts.
    2. Rotate the carousel 1/16 of a turn.
    3. Bring the carousel down on the mount pins.
  1. Close the load-lock door. See *Figure 7, Handle (Secured Position)*.
     1. Secure the top door handle. Do NOT overtighten.
     2. Use a crowbar on the lefthand side to help seal the door during pumpdown.

1. Pump Down the Load-lock Chamber
   1. Press Load lock AUTO PUMP. See *Figure 4, Load Lock Control Panel.*
2. The light for the Roughing Pump will come on.
3. Within 30 seconds the load lock handle will fall. This indicates that the door is sealed and the load-lock is pumping down.
4. After about 8 minutes the LOAD LOCK ROUGHING VALVE light will be OFF (automatically) and LOAD LOCK HIGH VAC VALVE light will be ON.



Load Lock High Vac Valve Light

Roughing Pump Button

Load Lock Roughing Valve Light

Figure 9, Load Lock Vacuum Panel

1. Check that the Load Lock Cryo-Pump is ON after the Load Lock opens.
   1. Turn off the roughing pump by pressing ROUGHING PUMP on the front control panel. See *Figure 9, Load Lock Vacuum Panel*.
   2. Set the timer for 5 minutes.
   3. Start the timer.
   4. Enter Coral run data while the load-lock is pumping down.
   5. When the 5 minutes are up, press GAUGE on the load-lock ion gauge to turn it on. See *Figure 2, Ion Gauges*.
   6. If the Main Chamber Ion Gauge is OFF, turn it ON by pressing GAUGE on the Chamber Ion gauge controller. See *Figure 2, Ion Gauges*.
2. Transfer Carousel to the Main Chamber
   1. Check to make sure the ion gauges are either both ON or both OFF or transfer gate valve will not open.

**CAUTION**


The Ion Gauges are extremely sensitive. They should only be turned on when the machine is under high vacuum and not sputtering.

* 1. Transfer to the main chamber may begin once the load lock IG pressure is less than 3 x 10-6 Torr.
  2. On the TMV computer (monitor) go to the MAIN MENU.
  3. Click “Substrate Transfer Control”. See *Figure 8, Main Menu*.



Figure , Main Menu

1. You will see Load lock, chamber and gate valve animated photographs here.
2. You can see if the gate valve is open or closed on this screen. Red = closed, Gray = open.




Transfer Gate Valve Open Button

Figure , Gate Valve Buttons

Transfer Gate Valve Close Button

* 1. Press the TRANSFER GATE VALVE OPEN on the front panel (manually). See *Figure 9, Gate Valve Buttons*.

1. You will hear a loud noise of the valve opening.
   * 1. Check on the monitor to see if the gate valve is open or closed.
2. **If gate valve is Open then proceed to next step otherwise contact Nanofab staff.**
   1. Press AUTO LOAD button on the control panel. See *Figure 10, Automatic Control Panel*.




Auto Load Button

Figure , Automatic Control Panel

Auto Unload Button

1. On the Computer Screen you will see animated pictures of carousel transferring to the main chamber.
2. The transfer process will take about 5 minutes.
3. The software will say “Transfer Arm is not at Home” during the transfer.
   * 1. Wait until the software says “Transfer Arm Is At Home”.
   1. When the Transfer Arm has returned home, press TRANSFER GATE VALVE CLOSE (red button) on the front panel. See *Figure 9, Gate Valve Buttons*.
4. You will hear loud noise of gate valve closing.
   1. Wait till Chamber IG pressure is below 9 x 10-7 Torr. See *Figure 2, Ion Gauges*.
5. This is called the base pressure.
   1. Write this value in the log book.
6. DC Sputter
   1. Turn OFF the Main Chamber ion gauge by pressing the gauge switch. See *Figure 2*.
7. There should be no reading on the IG display.
   1. Rotate the control mode key to the MANUAL position. See *Figure 3, Control Mode*.
   2. Go to the Main Menu.
   3. From the Main Menu on the computer click “Substrate Index Rotation Control”. See *Figure 8, Main Menu*.
   4. Click INDEX POWER. See *Figure 11, Substrate Index Rotation Control Window*.
   5. Wait for carousel to stop (on the picture).
   6. Check your sample position inside the main chamber using the flashlight.
   7. Note the chuck number on the carousel.
   8. Click NEXT INDEX to bring the samples in front of your desired metal target. See *Figure 18, TMV Target Positions.*
   9. Click on MAIN MENU.



Index Power

Next Index

Figure , Substrate Index Rotation Control Window

* 1. Click on PROCESS GAS CONTROL from the MAIN MENU.
     1. Click on CHAMBER BARATRON VALVE.
     2. Wait for 10 seconds until the indicator on the control panel lights up and stops blinking.
     3. Click on REMOTE PRESSURE SET POINT OFF.
     4. Click on the SET PRESSURE VALUE. See *Table 3, Sputter Parameters.*




Pressure Readout

Figure , MKS Pressure Controller

* + - 1. Type 20 to get 20 mTorr pressure.
    1. Click CHAMBER PROCESS GAS.
    2. Click INDIRECT.
    3. Click GAS-3 (Ar gas).

1. Flow rate of Argon is 150 sccm.
   * 1. If reactive sputtering with O2, click GAS 2.
        1. See staff for assistance switching from O2 to N2.
   1. Read the chamber pressure on MKS Pressure controller. See *Figure 12, MKS Pressure Controller*.
2. The reading should match the pressure you entered in step *12.8.4.1*.

Figure , Sample Rotation Selector



1 = Cr

2 = Ir or TiW

3 = Ti

4 = Au or Pt

* 1. Ensure that the AML (sample rotation power source) is on.
  2. Rotate the Sample Rotation Selector knob to the target location. See *Figure 13*.

1. Make sure to turn the Sample Rotation Selector knob to OFF before changing targets.
   * 1. Using the flashlight, check inside the chamber (through the window) to see if your samples are at desired location and the chuck is rotating.
   1. Click on MAIN MENU.
   2. From the main menu, click on CATHODE/SUPPLY SWITCHING.
      1. Ensure that the picture of the target you see on the monitor matches with the white board on TMV.




Figure , Power Supply/Cathode Selection Screen

1. Each target lists the power supply that connects to it.
   * 1. Click on the desired metal target.
     2. Make sure the connection is between DCG-1 / DCG-2 and the desired metal target and there is NO OTHER connection to any other metal target.
     3. If the connection is in RPG-1 or RPG-2 then click on the generator picture.
   1. Turn on the DCG (either DCG-1 or DGC-2). See *Figure 15.*



DCG 1

DCG 2

DCG Power Switches

MKS Pressure Controller

Figure , DCGs

Remote Enable Lights

* + 1. Wait a few moments till the generator is ready (wait for click).
    2. Make sure that the remote enable light is on for the power supply you turned on.

1. The power supply display will indicate that is it stable.
   1. Back on the computer monitor, click on the appropriate DCG tab.
   2. Click on the number next to TYPE IN SETPOINT.
      1. Adjust the number until you get the desired power for your sputtering. See *Table 1, Power Setpoints*. Check the display on the EVI power supply for actual power read out.

|  |  |  |
| --- | --- | --- |
| Table 1, Power Setpoints | | |
| Power (Watts) | Set Point DCG | Set Point RPG |
| 50 (max power for liftoff) | 0.94 | 0.42 |
| 100 | 1.92 | 0.92 |
| 150 | 2.88 | 1.4 |
| 200 | 3.86 | 1.9 |
| 250 | 4.85 | 2.38 |
| 300 (do not exceed 300 W) | 5.84 | 2.87 |

DO NOT exceed 300W

(50 W max for liftoff, photoresist will hard bake at higher power)

**CAUTION**

* 1. On the computer monitor, click on PLASMA to begin sputtering. See *Figure 16*.
  2. Look through the chamber porthole to see if plasma is formed.

1. The plasma is a purple or blue glow also the Voltage should drop when the plasma ignites.
   * 1. **If yes proceed to next step**, otherwise click on PLASMA again (to Stop) and go to section *15 Troubleshooting*.




Figure , Plasma Supply Control

Plasma Button


* 1. Pre-sputter for desired time (usually 2 min).

1. The Voltage should read between 200-700 Volts and the Current should read between 0.1-0.5 Amps.

|  |  |
| --- | --- |
| Table 2, Cathode Materials | |
| Target | Cathode Shutter # |
| Cr | 1 |
| TiW | 2 |
| Ir | 2 |
| Ti | 3 |
| Au | 4 |
| Pt | 4 |

* 1. Adjust Gas Pressure to desired level.
  2. Set timer to the desired sputter time. See **Error! Reference source not found.***.*
  3. Press CATHODE SHUTTER button on the front panel. See *Table 2, Cathode Materials*.

1. Shutter should be opened manually from the front panel and not from the software.
   1. Start the timer.
   2. Check inside the chamber from the window if shutter is open fully and if chuck is rotating.
      1. Any abnormalities need to written on the log book (in comment section).
   3. When the desired time has elapsed, press CATHODE SHUTTER on the panel manually to close the shutter. See *Figure 17*.
   4. If more samples on other chucks need to be sputtered with the same target then do the following:



Cathode Shutter Buttons

Figure , Cathode Shutter Control

* + 1. Turn OFF the Rotation Control.
    2. Click MAIN MENU.
    3. Click SUBSTRATE INDEX ROTATION CONTROL.
    4. Click on NEXT INDEX until your samples are in front of desired metal target.
    5. Repeat from Step *12.22*.
  1. When finished sputtering with the target, click on PLASMA (to stop).
  2. Turn off the DCG power. See *Figure 15*.
  3. If you need to sputter another layer of different metal then repeat the process from Step *12.12*.
  4. When Sputtering is completely finished, click on MAIN MENU.
  5. Click on PROCESS GAS CONTROL.
     1. Click on GAS 3 to stop gas flow.
     2. Click on INDIRECT to stop.
     3. Click on CHAMBER PROCESS GAS.
     4. Click on REMOTE PRESSURE SET POINT ON.
     5. Click on CHAMBER BARATRON VALVE.
  6. Click on MAIN MENU.
  7. Rotate the Sample Rotation Selector knob to OFF. See *Figure 13*.

1. Transfer of Carousel back to Load-Lock
   1. Make sure the DCG power is off, the gas pressure is zero, and the Sample Rotation Selector is OFF.
   2. Press Gauge to turn ON the Main Chamber ion gauge.
   3. Make sure that the Load Lock Ion Gauge is also ON.
   4. On the computer go to the MAIN MENU click “Substrate Transfer Control”.
      1. The computer screen should show the transfer gate valve closed (red triangles).
   5. Press TRANSFER GATE VALVE OPEN on the front panel (manually).
2. You will hear loud noise of valve opening.
   1. Check on the monitor to see if gate valve is open or closed (red triangles are gone).
      1. If gate valve is OPEN then proceed to next step otherwise contact fab staff.
   2. Turn the key to AUTO mode on the control panel.
   3. Press AUTOMATIC UNLOAD on the control panel.
3. You will see animated pictures of transfer of carousel from main chamber to load-lock.
4. The transfer process will take about 5 minutes.
   * 1. Wait until the software says “Tranfer Arm is at Home”.
     2. Contact staff if transfer fails.
   1. Press TRANSFER GATE VALVE CLOSE on the front panel.
5. You will hear loud noise of gate valve closing.
6. Unload Samples from the Load-lock
   1. Turn off the load lock ion gauge.
      1. Make sure that there is NO reading on load lock ion gauge.
   2. Press AUTO VENT on the load-lock panel.
7. You will hear loud noise of air rushing to load lock.
   1. Wait for 1-2 min and Load-lock door will OPEN automatically.
   2. Press AUTO VENT on load-lock panel to stop gas flow.
   3. Unload the carousel. Ask for assistance if needed.
      1. Remove samples.
   4. Mount the carousel in the load-lock chamber. Ask for assistance if needed.
   5. Put the bottom load-lock locking handle in a horizontal position.
   6. Turn on N2 on wall behind TMV.
   7. Press Auto Pump on the load lock control panel.
      1. When the roughing pump light comes on, start the roughing pump.
   8. Once the load lock high-vac valve has opened as indicated on the control panel, turn off the roughing pump.
   9. Turn off the N2 gas.
   10. Turn off the load lock Auto Pump
   11. Make sure your run has been entered into Coral.
   12. Disable the TMV system in Coral.
8. Troubleshooting

Plasma does not Ignite

* + 1. Go to the Process Gas Control page.
    2. Click on the pressure set point.
    3. Increase the set point.
    4. Click Ok.
    5. Go back to the DCG page.
    6. Click Plasma to start sputtering.
    7. Check the Voltage, it should drop to within the 200-700V range when the plasma ignites.
    8. If the plasma still does not ignite, try opening the shutter.
    9. If the plasma does ignite, close the shutter, if open.
    10. Go back to the Process Gas Control page to correct the gas pressure setpoint.

1. Process Notes

TMV Parameters

|  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Table 3, Sputter Parameters | | | | | | |
| Shutter # | Metal | Power  (Watts) | Pressure  (mTorr) | Time  (min) | Thickness  (nm) | Dep Rate  (nm/min) |
| 1 | Cr | 100 | 10 | 10 | 101 | 10 |
| 2 | TiW | 45 | 10 | 30 | 191 | 6.4 |
| 2 | Ir | 100 | 10 | 10 | 80 | 8 |
| 3 | Ti | 90 | 10 | 10 | 70 | 7 |
| 4 | Au | 45 | 10 | 10 | 190 | 19 |
| 4 | Pt | 90 | 10 | 20 | 336 | 16 |

TMV Target Positions

Source 3 Shutter

Back

**Ti**

DCG2

DCG1

**Cr**

Front

Source 1 Shutter

Source 2 Shutter

DCG1

Source 4 Shutter

DCG2

Top

**TiW**

DCG1

Top

**Au**

DCG2

Bottom

**Pt**

DCG2

Bottom

**Ir**

DCG1

Figure , TMV Target Positions

TMV DOOR

1. Revision History

|  |  |  |  |
| --- | --- | --- | --- |
| **Rev** | **Date** | **Originator** | **Description of Changes** |
| 1 | 10 May 2010 | L. Williams | Initial release. |
| 2 | 20 May 2010 | Sam Bell | Added figures and tables. |
| 3 | 28 Dec 2010 | Sam Bell | Deleted fig. 9, added text |
| 4 | 4 Apr 2011 | Sam Bell | Changed rotation control figure and procedure pp 11, 14, 15. Updated Table 3, pp 16 |