4 Point Probe SOP

- Scope

  - This document provides the procedure for measuring and calculating sheet resistance and resistivity with the 4 point probe.

- Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	2

3.1	Referenced within this Document	2

3.2	External Documents	2

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Measurement Procedures	2

7.	Calculation Procedures	3

8.	Revision History	3

Figure 1, 4 Point Probe Controls	2

- Reference Documents

- Referenced within this Document

    - None

- External Documents

    - None

- Equipment and/or Materials

  - 4 Point Probe

  - Wafer/Sample

- Safety

  - Follow all Nanofab safety procedures.

- Measurement Procedures

  - Turn on the Probe.

    - Turn on the power switch. See Error: Reference source not found.

  - Load the Sample.

Voltage Knob

Current Knob

Contact Lever

Power Switch

Figure , 4 Point Probe Controls

  - Turn the current down to 1 mA using the knob on the base of the probe station. See Error: Reference source not found.

  - Turn the contact lever to the left to contact the substrate with the probe tips. See Error: Reference source not found.

  - Turn the voltage and current knobs until you get a reading that is not 1 or 0. See Error: Reference source not found.

  - Note the measured voltage and the current you are applying.

  - Use the equations below to calculate sheet resistance and resistivity.

- Calculation Procedures

  - On the desktop of the computer next to the 4 point probe, find the Excel sheet named Resistivity Calculator and plug in your values.

  - Sheet Resistance (Rs) = resistivity/thickness of conductive layer(s) = 4.53*V/I

  - Where V is the voltage measured by the 4 point probe and I is the current applied to the probes.

  - For Example:

  - 0.2 um film of gold on a glass substrate measures 2.05 mV on the 10 mV scale with a current setting of 44.5 mA.

  - The sheet resistance would be:

  - Rs (Ohm/square) = 4.53*V/I = 4.53*.00205V/.0445A = 0.209 Ohm/square.

  - And the resistivity would be:

  - Resistivity (Ohm-cm) = Rs*t = .209 Ohm/sq *0.2 um *.0001 cm/um = .00000418 Ohm-cm = 4.18 uOhm-cm.

  - Another Example:

  - 2” N-type wafer with a specified resistivity of 1-10 Ohm-cm and a thickness of 10-12 mils (250-300 um).

  - The voltage measured is 123.4 mV on the 100 mV scale and the current is 4.53 mA.

  - Rs = 4.53 * V/I = 4.53 * 0.1234V/.00453A = 123.4 Ohm/sq.

  - Resistivity = Rs*t = 123.4 Ohm/sq * 275 um * .0001cm/um = 3.4 Ohm-cm, which is in the specified range of 1-10 Ohm-cm.

- Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 02 June 2010 | Sam Bell |  |