D150 GRC SOP

Scope

This document provides the operating procedure for the D150 Dual Gas Reactor Column.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	1

3.1	Referenced within this Document	1

3.2	External Documents	1

4.	Equipment and/or Materials	1

5.	Safety	2

6.	Setup Procedures	2

7.	Operating Procedures	2

7.2	Automatic preheat and changeover	3

7.3	Manual preheat and changeover	3

7.4	Shutdown	3

7.5	Figures	4

7.6	Warnings	7

8.	Process Notes	7

8.1	Fault indications	7

8.2	GRC Characteristics	8

8.3	Process Summary	8

9.	Revision History	9

Figure 1, GRC Controller	4

Figure 2, GRC Cabinet	6

Table 1, Controller Components	5

Table 2, Cabinet Components	7

Table 3, Fault Conditions	8

Reference Documents

Referenced within this Document

None

External Documents

D150 Dual Gas Reactor Column Instruction Manual, Edwards High Vacuum International.

Equipment and/or Materials

D150 Dual GRC

N2 supply

Compressed air supply

Safety

Follow all Nanofab safety procedures.

The D150 has many safety features, including interlocks and protective enclosures to prevent accidental contact with hot components, valve interlocks to prevent inadvertent valve operation, overheat protection, and overpressure protection.

Setup Procedures

Ensure the D150 has N2 and compressed air supplies.

Operating Procedures

Initial Start-up

Ensure that only purge gas is flowing through the pumping system.

Refer to Figure 1. Ensure that the controls on the control unit are in the correct position:

Electrical supply isolator (9) is in the OFF position

Safety interlock key (10) is in the CARTRIDGE CHANGE position

Button lock key (13) is in the LOCKED position

Refer to Figure 2. Ensure that the electrical supply and control connectors (2 and 3 or 8 and 9) are fitted to one of the heater units (7 and 4).

Ensure that the off-line purge and pressure-switch purge control valves (Fig. 1, 1 &2) are fully closed.

Open end-point isolation valves.

Turn on the N2 supply and compressed air supply to the correct pressures: N2 80 psi, Air 40 psi.

Switch on the electrical supply isolator (Fig. 1, #9).

Check that the pump start valve LED is on to indicate that the start-up valve is open.

Turn the safety interlock key to the UNLOCKED (fully clockwise) position; then immediately turn the off-line purge flow control valve anticlockwise so that a flow of 1 cm/h is shown on the flow meter.

Check that the leak test fail LED flashes to indicate that leak test is in progress.

After 70 seconds, the leak test fail LED should go off to indicate that the leak test was successful. If the LED goes off, continue at 7.1.13 otherwise continue at 7.1.12.

If the leak test LED goes permanently on (that is, the leak test failed):

Turn the safety interlock key to the LOCKED position.

Check that the cartridge inlet and outlet pipeline connections are correctly fitted.

Remove the electrical supply connector from the heater unit, wait 5 seconds, then refit the connector.

Turn the safety interlock key to the UNLOCKED position to leak test the cartridge again. If the leak test continues to fail, the process pipeline connections may be incorrectly fitted, or the cartridge may not be leak tight and must be replaced.

Close the enclosure door for the available GRC.

Ensure that the button lock key is in the UNLOCKED position.

Press and hold the changeover button for 10 seconds, then look at the status display LEDs:

The corresponding on-line LED for the GRC goes on to indicate that the heater is starting to heat up.

After 10 seconds, the pump start LED goes off to indicate that the pump start valve has closed.

The corresponding changeover LED and the GRC outlet isolation valve LED for the on-line GRC go on to indicate that process gases now pass through the cartridge in the on-line GRC.

The temperature LED and the level 1 warning LED on the status display go on to indicate the cartridge in the on-line GRC has not yet reached operating temperature.

The on-line GRC will take ~5 hours to reach operating temperature.

Automatic preheat and changeover

After Initial Start-up the only operator action required is to change cartridges when necessary.

The D150 will automatically start to preheat the available GRC when required and will automatically changeover.

Manual preheat and changeover

Manually select preheat

Ensure the button lock key is in the UNLOCKED position.

Press and release the preheat button. The preheat LED will go on to indicate that the available cartridge has started to heat up.

To cancel preheat, press and hold the preheat button for five seconds.

Manually select changeover

Ensure the button lock key is in the UNLOCKED position.

Press and release the changeover button. The available GRC will come on-line in 5 hours, when it reaches operating temperature.

To changeover immediately, press and hold the changeover button for 10 seconds.

To cancel changeover, press and hold changeover button for 5 seconds.

Shutdown

Ensure that the system has been fully purged.

Switch off the electrical supply isolator and leave the D150 to cool for at least 24 hours.

Open the door of the off-line GRC, then push the rod on the door lock and open the other door.

Remove the electrical supply and control connectors from the heater units (Fig. 2).

Switch on the electrical supply isolator and check that the pump start LED goes on to indicate that the pump start valve is open.

If required, remove the heater units.

Switch off the electrical supply isolator, then turn off the external electrical supply and the compressed air and nitrogen supplies.

Figures

Figure 1, GRC Controller

| Table 1, Controller Components |  |
| --- | --- |
| - # | - Component |
| - 1 | - Off line purge flow meter and control valve |
| - 2 | - Pressure switch purge flow meter and control valve |
| - 3 | - TMS temperature controllers |
| - 4 | - Upper cartridge temp. controller (GRC 1) |
| - 5 | - Lower cartridge temp. controller (GRC 1) |
| - 6 | - Status display |
| - 7 | - Upper cartridge temp. controller (GRC 2) |
| - 8 | - Lower cartridge temp. controller (GRC 2) |
| - 9 | - Electrical supply isolator (Power switch) |
| - 10 | - Safety interlock key |
| - 11 | - Changeover button |
| - 12 | - Preheat start button |
| - 13 | - Button lock key |
| - A | - Control unit |
| - B | - Lower control panel |

Figure 2, GRC Cabinet

| Table 2, Cabinet Components |  |
| --- | --- |
| - # | - Component |
| - 1 | - Cartridge outlet clamp (GRC2) |
| - 2 | - Electrical supply connector (GRC2) |
| - 3 | - Electrical control connector (GRC2) |
| - 4 | - Heater unit (GRC2) |
| - 5 | - Cartridge inlet clamp (GRC2) |
| - 6 | - Cartridge inlet clamp (GRC1) |
| - 7 | - Heater unit (GRC1) |
| - 8 | - Electrical control connector (GRC1) |
| - 9 | - Electrical supply connector (GRC1) |
| - 10 | - Cartridge outlet clamp (GRC1) |

Warnings

|  | WARNING If the flow of untreated process gas into your exhaust-extraction system may result in a hazardous situation, you must incorporate suitable safety precautions in your control system. Do not operate the D150 with gas flow (through a GRC cartridge) less than 5 slm or greater than 60 slm. If you do, the D150 may not operate correctly and your exhaust-extraction system may be damaged. If a warning or alarm LED goes on, take the necessary operator action as soon as possible. If you do not, untreated process gas may flow directly into the exhaust-extraction system. |  |
| --- | --- | --- |

Process Notes

Fault indications

In normal operating conditions, once one GRC is on-line and the other GRC has a new cartridge and has been successfully leak tested, only green LEDs on the status display should be on.

f one of the warning or alarm LEDs on the status display goes on, you must determine the cause and take necessary operator action. The conditions which will cause the warning and alarm LEDs to go on are described in Table 4. The priorities of the warning and alarm LEDs are as follows:

If the warning level 1 LED is on, operator action is required within 24 hours.

If the warning level 2 LED is on, operator action is required as soon as possible.

The alarm LED goes on when the pump start valve is open (that is, process gas passes directly from the D150 inlet to the D150 outlet) and should only go on during initial start-up. If the alarm LED goes on at any other time, you must shut down the system immediately.

| Table 3, Fault Conditions |  |  |  |
| --- | --- | --- | --- |
| - Fault type | - Fault condition | - Open contacts | - Closed contacts |
| - Warning level 1 | - End-point detector flow fault - Pressure warning - Leak test failed - Cartridge awaiting change | - Pins 4 and 9 | - Pins 4 and 5 |
| - Warning level 2 | - High Pressure - Circuit breaker tripped - Compressed-air supply failure - Temperature fault | - Pins 7 and 3 | - Pins 7 and 8 |
| - Alarm | - The pump start valve is open | - Pins 1 and 6 | - Pins 1 and 2 |

GRC Characteristics

If there is an ‘available’ GRC, automatic changeover is initiated by one of three events:

There is an increase in back-pressure in the on-line GRC inlet pipelines.

The on-line cartridge has been in use for its maximum predicted life: default = 60 days.

The end point detector for the on-line GRC indicates higher concentration of process gas in the cartridge outlet pipeline (breakthrough).

Process Summary

The automatic changeover sequence is as follows:

The available GRC is heated up for five hours (this is known as preheat).

After the five hour preheat, the on-line GRC goes off-line and the available GRC becomes the on-line GRC.

The off-line GRC is purged with nitrogen and cooled for 24 hours.

After the 24 hours, the off-line GRC is in the await change state.

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 21 Dec 2009 | Sam Bell |  |