Heidelberg Pattern Generator SOP

Scope

This document provides the operating procedures for the Heidelberg Pattern Generator with Version 3.12.5 software.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	2

3.1	Referenced within this Document	2

3.2	External Documents	2

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Pattern Design and Conversion	2

6.1	Design Data	2

6.2	General Design Rules	3

6.3	DXF Specific Design Rules (AutoCAD)	3

6.4	CIF Specific Design Rules	3

6.5	BMP Specific Design Rules	4

6.6	GDSII Specific Design Rules	4

6.7	Convert to CIF format	4

7.	µPG Normal Operations	4

7.1	Reserve and Enable µPG in Coral	4

7.2	Setup PC	5

7.3	Load Design	5

7.4	Set PG Parameters	6

7.5	Load Mask	7

7.6	Manual Alignment	10

7.7	Begin Expose	11

7.8	Unload Mask	12

8.	Troubleshooting Problems	13

9.	Alignment Options	13

10.	Revision History	15

Reference Documents

Referenced within this Document

None

External Documents

None

Equipment and/or Materials

Mask

USB Drive

Safety

Follow all Nanofab safety procedures.

Pattern Design and Conversion

Design Data

The μPG 101 can expose designs created in one of the following formats

- CIF: Caltech Intermediate Form, easy-to-use language for direct definition of

structures in a text file  (recommended format)

- DXF: Standard AutoCAD format

- BMP: Standard MS-Windows® Bitmap format. This is the only format

useable for grayscale exposures (3D resist structuring

- GDSII: Standard format for dense binary (2 dimensional) designs

NOTE: To get the correct results, certain rules have to be followed when creating the design.

General Design Rules

Maximum design size is 100 mm x 100 mm.

Use polygons or circles that enclose areas, these may overlap and will be combined for exposure with an OR function.

Defined areas will be exposed by default (you can select darkfield or lightfield later in the μPG program).

Do not use special characters (spaces, punctuation marks, umlauts ..) in the design name.

It is recommended to center the design (0, 0 at center of pattern), otherwise the μPG program will calculate offsets using max and min x and y and center the design.

DXF Specific Design Rules (AutoCAD)

Use a 100% AutoCAD compatible editor.

The design has to be created in metric units.

Do not use single lines without width.

All polylines must be closed (use close command for last segment). Keep in mind interiors of polyline structures will be filled out completely, covering any structures within.

Do not cross polylines, or create doubled vertices (vertices with no distance between them). This will create data error.

As far as possible, only use the structure types “polyline”, “circle”, and “text”.

Only one text font is provided. This will replace any other font used in the design creation. The only supported text attributes are “rotated”, “mirrored”, and “scaled”.

Try to use as few layers as possible. All exported layers will be merged via an OR operation.

Do not use special characters in the layer names.

Polylines with width must not change their width (no tapered lines).

When inserting blocks, the same scaling has to be used for x and y.

External blocks are not supported.

- NOTE: Zoom Extents will find any unexpected features outside normal design viewing window.

Save file as .dxf, use LinkCAD to convert to .cif., see following section Convert to CIF format

CIF Specific Design Rules

Use only one layer. If several layers exist in a design file, only the layer with the most references will be written.

Maximum definition or reference depth is 50.

The subroutine with the highest number definition will be taken as main routine.

BMP Specific Design Rules

Maximum image file size is 4 GB.

For grayscale exposures, use 8-bit format. The 256 gray values will be interpolated down to 100 gray values.

Create the design in a 90° counter-clockwise rotated version, as the design will be rotated 90° clockwise during data processing.

GDSII Specific Design Rules

No Inclusions: Do not include other GDII files or text libraries, they will be ignored.

No nodes: Node statements in GDSII are ignored.

Convert to CIF format

Start LinkCad software on cad station.

Select import format (i.e. DXF) from dropdown list.

Select Export format CIF, then click on next.

Check Units and scaling factor and adjust for input drawing, try using defaults for ARC Precision and Layer Options, if text is in design remove check to ignore text.

Click Next

Try using defaults for CIF Export Options, click Next

Select your import file for conversion in the browsing window, click Open

Note error Log Messages if present and click Next (not shown if no errors)

Check correct Layers included in File Structure screen, click Next

Click Save to use same name with CIF extension, then Quit

Find the saved .CIF file and double click to launch in viewer

Confirm conversion was successful, note origin, design size and offsets

Exit .CIF viewer

µPG Normal Operations

Reserve and Enable µPG in Coral

Reserve time for the Pattern Generator on the computer reservation system.

Enable the PG in Coral before beginning operation.

Keep track of the following data to input into the Coral log:

| 1. | Substrate |
| --- | --- |
| 2. | Photoresist Type |
| 3. | Design file name |
| 4. | File type |
| 5. | Laser power |
| 6 | Pixel energy |
| 7. | Size-X |
| 8. | Size-Y |
| 9. | Inverted Y/N |
| 10. | X offset |
| 11. | Y offset |
| 12. | Number of Stripes |
| 13. | Alignment No/Auto/Manual |
| 14. | Total exposure time |

Setup PC

Go to the µPG PC

Click on labuser icon to get to the desk top

Copy design file onto the PC

Insert flash drive

Go to the Utah Users folder on the desktop

Find your directory or create a new folder titled ‘your name’

Copy your design file(s) into this folder

Start Exposure Wizard. See Figure 1, Exposure Wizard Icon

Figure 1, Exposure Wizard Icon

Wait for system initialization, lid must be closed for initialization or stage movement

Make sure the box for Show Advanced Options is checked and press Next

- NOTE: Settings from previous exposure will be defaults without Advanced Options and not in standard selection menus

Load Design

Press Load Design, See Figure 2, Image Preview

Select your file

Enter units – default 1000 for DFX ( for CIF units = 1)

Wait for preview image to process

Click on image to check it and close preview after checking

If not correct, either return to Load Design step and load the correct design or select EXIT to quit Wizard

Make sure size x, y and offset x, y look correct

Note size and offset for Log

Click Next

Default exposes dark area

Check ‘inverted’ box to expose white space

Figure 2, Image Preview

Set PG Parameters

Check and set Options: see Figure 3, Check Options – Expose Window

Power in mW, see current recommendation tag on μPG101

Exposure level %, see current recommendation tag on μPG101

Energy Mode:  1x1 (default for typical masks and resist)

Inverted check on = Light field (opaque where features are drawn) or Inverted check off (default) = dark field (transparent where features are drawn), see Figure 2, Image Preview

- NOTE: RIE or Wet Etch processes typically use light field mask to protect feature layer with PR. Lift-off processes typically use dark field mask for layer deposition into open areas

Enter size of frame to expose extending outside design for inverted exposure – 0 to 10000 μm (larger number will increase write time)

If the design is NOT centered and “no alignment” is planned, check Automatic Centering and if check box is active

Figure 3, Check Options – Expose Window

Click Show Control Panel button

Check for non-centered designs

Figure 3, Check Options Window

14

40

Load Mask

Click “To Un/Load” button in Stage Controls, see Figure 6, Controls Window

Load Mask

Open lid

Place a 5 x5” mask against the 3 centering pins in the stage, 1 on the front edge and 2 on the left edge

Contact staff for using substrates 1” or smaller, Writehead cannot auto focus with 1” or smaller size

Or for other size substrates visually center on the stage

If the pins must be removed to accommodate a larger substrate, contact staff

- NOTE: Pins are an aid for centering the mask on the stage but not required

Figure 4, Placement & Vacuum Select

and 5”:

(except 5” requires no pins)

White Knobs

Vacuum Knob

Figure 5, μPG Chuck

OFF

ON

Check to make sure the mask is photoresist side up

Turn on vacuum knob, see Figure 5, μPG Chuck

Adjust white knobs for mask size if needed, see Figure 4, Placement & Vacuum Select

PULL OUT PINS *IMPORTANT*, if the pins protrude above substrate surface (for non-standard or thin substrates)

|  | WARNING Severe damage to the Writehead could occur if pins are above mask surface. |  |
| --- | --- | --- |

Shut lid

Click “To Center” on the Stage Control, the stage will move to stage center under Writehead

Click “Focus” on the Writehead Control, the Writehead will move to focus and expose height above substrate

- NOTE: Focus is complete when a green rectangle is present above Focus button, do not proceed until green rectangle is present, see Figure 6, Controls Window

Click “Close” in the Controls window

Click “Next” in the Check Options window

Figure 4, Expose Window

LEAVE UNCHECKED

Figure 6, Controls Window

Manual Alignment

Default is no alignment – design centered to stage center

Confirm Offsets are appropriate for design, see Figure 7, Manual Alignment

Figure 7, Manual Alignment

Contact staff for use of Manual Offsets and Alignment Options , some instructions are provided in last section, Alignment Options

Click NEXT

- NOTE: For non-BMP designs, exposure is rotated in relation to front of Heidelberg, See Figure 8, Stage Coordinates

Figure 8, Stage Coodinates

Begin Expose

Check parameters, see Figure 9, Expose

Check Auto Unload after Exposure

Click EXPOSE

Figure 9, Expose

Unload Mask

Unload mask when exposure is complete

Turn off vacuum

Wait 10 seconds

Unload mask

Replace pins at 5” mask locations if removed

Click OK in pop-up window

Click Exit Wizard

Enter information into Coral

Log out of μPG PC

Troubleshooting Problems

Terminating the Program during Normal Operation

Click the blue Abort! Button

Alignment Options

Manual Alignment Options, contact staff

Option - Find Plate Center

Select check box fast mode and start to auto locate substrate edges and calculate offsets to center design on substrate, see (insert ref)

Select accept to apply calculated offsets

Figure 10, Find Plate Center

Select exit or cancel to exit window without offsets

Figure 11, Setup Automatic Alignment

Option, Advanced Alignment -

Contact staff to assist using Advanced alignment, see (insert ref)

For POS enter design position of mark center in µm from design center, required for either Manual or Auto

Only use Load or Create to use mark image for Auto align

Check Mark 1 & Mark 2 to locate marks and calculate offsets and rotation

Check Mark 1 only to locate one mark and calculate offset

Select Execute alignment and use Stage Controls in Controls window to move to mark location

    - For Manual, select POI and move crosshair to center of mark

Option Manual Offsets

Select Manual Offsets

Enter offset value in pop-up windows for x and y offsets

Option Manual Rotation

Select Manual Rotation

Enter offset value in pop-up windows for rotation in degrees

Click Next after completing any desired positioning selections to exit Manual Alignment Window

Figure 12, Camera View Window for Positioning & Alignment

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 18 Oct 2011 | Sam Bell |  |
| 2 | 8 Aug 2012 | Steve Pritchett | Modify for software version on new Heidelberg |