Keithley Test Station SOP

Scope

This document provides procedures for IV and resistance, CV, Diode, and Transistor measurement using the Keithley test station.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	1

3.1	Referenced within this Document	1

3.2	External Documents	1

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Setup Procedures	2

7.	IV and Resistance Measurement Procedures	3

8.	CV Measurement Procedures	5

9.	Transistor Measurement Procedures	7

10.	Diode Measurement Procedures	7

11.	Viewing and Saving Results	8

12.	Graph Examples	9

13.	Revision History	11

Figure 1, SMU Connections	2

Figure 2, Sample Loading	3

Figure 3, KITE IV Window	3

Figure 4, IV Parameters	4

Figure 5, KITE CV Window	5

Figure 6, Capacitance CV Meters	5

Figure 7, KCON Window	6

Figure 8, Transistor Window	7

Figure 9, Resistor Graph	9

Figure 10, CV Graph	9

Figure 11, Diode Graph	10

Figure 12, Transistor Graph	10

Reference Documents

Referenced within this Document

None

External Documents

Keithley Application Notes for C-V Measurements

Equipment and/or Materials

Keithley 4200-SCS

Probe Station

Sample

Safety

Follow all Nanofab safety procedures.

Setup Procedures

Double click the KITE icon on the desktop.

Click yes.

Check to make sure the program list as shown in Figure 3 says default at the top, if not:

Go to File > Open Project.

Choose the default.kpr  file from S4200/kiuser/projects/default.

Check SMU connections.

Check the numbers of the SMU cables that the probes you are using are connected to.

Figure 1, SMU Connections

Numbers on this side correspond to the KITE Software SMU numbers.

Load sample. See Figure 2, Sample Loading.

Manually move probes into place gently.

Use knobs to lower probes until tips contact measurement points on the sample.

Figure 2, Sample Loading

Probe Unit

Knobs

Probe Tip

Select SMU number

Click to run test

Click this bar to set test parameters

Program List

Figure 3, KITE IV Window

IV and Resistance Measurement Procedures

Select res2t under 2-wireresistor from the program list. See Figure 3, KITE IV Window.

Select SMU number.

Click on the SMU window drop down menu. See Figure 3.

Select correct SMU number. See Figure 1, SMU Connections.

Set voltage and step values.

Click on the SMU window bar. See Figure 3.

Select the type of device (generally one Sweep Voltage and one Common).

For the Sweep Voltage device set the voltage and step values. See Figure 4, IV Parameters.

Click OK.

Figure 4, IV Parameters

Press green arrow to run test. See Figure 3.

See Section 11 for viewing and saving results. See Section 12.1 for an example Resistor graph.

Figure 5, KITE CV Window

CV Measurement Procedures

Turn on capacitance meter, voltage source, and CV analyzer. See Figure 6, Capacitance CV Meters.

Figure 6, Capacitance CV Meters

Exit KITE

Open KCON

Go to Tools

Select Add External Instrument>Capacitance Meter. See Figure 7, KCON Window.

Connect instruments 590, 595, and 82

Close KCON

Open KITE

Select cv under 4terminal-n-fet from the program list. See Figure 5, KITE CV Window.

Set all parameters as needed.

Click green arrow to begin test.

NOTE:	See the Keithley Application Notes for C-V Measurements for a more detailed description of the CV measurement process.

See Section 11 for viewing and saving results. See Section 12.2 for an example CV graph.

Figure 7, KCON Window

Transistor Measurement Procedures

Click on vg-ic under 3terminal-n-fet on the program list.

Set parameters for each terminal. See Figure 8, Transistor Window.

Click on the correct SMU header for each terminal.

Set SMU parameters according to its function.

Figure 8, Transistor Window

Click the green arrow to start test.

See Section 11 for viewing and saving results. See Section 12.4 for an example Transistor graph.

Diode Measurement Procedures

Select the ‘vfd’ (voltage forward bias) or ‘vrd’ (voltage reverse bias) in the left-side frame and a diode symbol will appear in the right-side frame.

For forward bias testing, select the ‘vfd’. Similar to the resistor, there will be 2 dialogue boxes. The Anode box should read SMU1, while Cathode box should read SMU2. Select the Anode Force Measure  button to setup the SMU1 test parameters.

You are now ready to perform diode measurement. Click the RUN icon to start measurement. Click on the GRAPH tab to view the I-V plot. Save your data as spelled out in resistor testing.

For reverse bias testing, select the ‘vrd’. Similar to the resistor, there will be 2 dialogue boxes. The Anode box should read SMU1, while Cathode box should read SMU2. Select the Anode Force Measure button to setup the SMU1 test parameters. Apply a START VOLT of 0V and STOPVOLT of -20V, 0.1V step by changing the values in the appropriate boxes. If no reverse breakdown is observed, increase the STOP VOLT using -20V intervals (e.g. -40V, -60V …, -140V). Increase voltage step size accordingly.

See Section 11 for viewing and saving results. See Section 12.3 for an example Diode graph.

Viewing and Saving Results

Click on Graph tab.

Set up graph parameters by right clicking.

Choose Define Graph

Click on Autoscale to view graph.

To save results:

Click on the Sheet tab.

Press Save As to save results in excel format.

See Section 12 for Graph examples.

Graph Examples

Resistor

Figure 9, Resistor Graph

CV

Figure 10, CV Graph

Diode

Figure 11, Diode Graph

Transistor

Figure 12, Transistor Graph

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 10 Aug 2010 | Sam Bell |  |