OAI 800 Backside Aligner

- Scope

  - This document provides the procedure of operating the OAI 800 Backside Aligner

- Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	2

3.1	Referenced within this Document	2

3.2	External Documents	2

4.	Equipment and/or Materials	2

5.	Safety	2

6.	Setup Procedures	2

7.	Alignment Procedure	3

7.2	Align Wafer	4

7.3	Expose Wafer	5

8.	Revision History	6

Figure 1: Objectives/Camera Assembly	2

Figure 2: Process Setting	3

Figure 3: Level Screen	3

Figure 4: Digital Video Image Generator	4

Figure 5: Optics	4

Figure 6: Wafer Alignment	5

- Reference Documents

- Referenced within this Document

- None

- External Documents

- None

- Equipment and/or Materials

  - OAI 800 Backside Aligner

  - Mask

  - Wafer/Sample

- Safety

  - Follow all Nanofab safety procedures.

  - Avoid exposure to UV output beam.

- Setup Procedures

  - Turn on Machine

  - If LCD screen is on, skip to 6.2

  - Ensure that the red switch on the power bar located at the rear of the tool is on and glowing.

  - Ensure that the breaker switch located at the rear of the tool is switched on.

Insert Mask

Figure 1: Objectives/Camera Assembly

  - Move the objectives/camera assembly to the back by pressing and holding the two control buttons on the arm and pushing backwards. See Figure 1: Objectives/Camera Assembly.

  - Loosen the screws holding the mask plate down and gently remove the plate by lifting the handles and sliding the mask plate off the aligner stage.

  - Be careful, the mask plate is HEAVY and can damage the camera assembly!

  - Hold the mask plate upside down and pull the silver handle of the mechanical mask clamp to release it.

  - Place your mask on the plate making sure it is centered by pushing it against the pins from the left lower corner.

  - The photoresist side should face up and the shiny side should face the mask holder.

  - Push the black lever on the mechanical mask clamp to clamp the wafer down.

  - Press MANUAL TEST > MASK VACUUM on the LCD screen to hold the mask. The button should turn black.

  - If you hear hissing, turn MASK VACUUM off and reseat the mask. Avoid riding on the 3 pins.

  - Load the mask/mask plate back onto the aligner stage by sliding it into position over the wafer chuck.

  - Be careful not to hit the camera assembly while moving the mask plate in!

  - Screw down the mask plate FIRMLY.

  - All 3 screws should be tight.

  - Mask sure mask holder handles are completely down or they can damage the microscope objectives.

  - Set Exposure and Gap

  - From the MAIN MENU press PROCESS SETTING > EXPOSURE TIME and enter the desired exposure time in seconds. See Figure 2: Process Setting.

  - Select the desired side alignment from PROCESS MODE in the same menu using the arrows                (I MPRINT, IR BACK, BACK, or FRONT).

From PROCESS SETTING press ADDITIONAL SETTING and enter the desired gap in microns (Minimum of 250).

Figure 2: Process Setting

  - If multiple exposures are desired on the same wafer, press RETURN and enter NUMBER OF CYCLES and OFFTIME IN SECONDS. See Figure 2.

  - Press LEVEL SCREEN.

- Alignment Procedure

  - Load Wafer

  - Pull the wafer chuck out by gently pulling the stainless steel handle in the front.

  - Be careful not to pull the wafer chuck too hard that it slams into the stopper!

  - Place your wafer on the chuck with the wafer flat at the bottom and photoresist side facing up.

From the LEVEL SCREEN select SUBSTRATE VACUUM to hold the wafer to the chuck. The button should turn black. See Figure 3: Level Screen.

Figure 3: Level Screen

  - Ensure the Z-Chuck electromotor is set to ~20 mA. (located on the bottom right corner of the aligner)

  - Gently and slowly push the wafer chuck back in.

  - Press LEVEL from LEVEL SCREEN.

Align Wafer

Figure 4: Digital Video Image Generator

  - Mask Alignment:

  - On each DIGITAL VIDEO IMAGE GENERATOR press 1 ONLY to display just the top camera display. See Figure 4: Digital Video Image Generator.

  - If the monitor displays HOLD in the top left of the screen whenever a button is pressed, press and hold LOCK on the DIGITAL VIDEO IMAGE GENERATOR for 3-4 seconds.

  - Turn on camera illumination by toggling the switch on the bottom left of the machine to TOP OPTICS. See Figure 5: Optics.

Adjust the illumination knobs as needed.

Figure 5: Optics

  - Center the object/camera field of view to the center of the mask by pressing the two buttons on the back of the arm and moving the camera assembly.

  - The top button is for X axis movement, the bottom button is for Y axis.

  - Move the individual objectives along the X axis by adjusting the large silver knobs on the side of each objective.

  - Adjust the theta axis by adjusting the micrometer behind the objectives.

  - Adjust the magnification and focus by rotating the adjustment wheels on the microscope objective.

  - The top adjustment knob is magnification, the bottom is focus.

  - Wafer Alignment:

  - Adjust the X axis by adjusting the micrometer on the front center of the wafer chuck.

  - Adjust the Y axis by adjusting the micrometer on the right side of the wafer chuck.

Adjust the theta axis by adjusting the micrometer on the front left of the wafer chuck. See Figure 6: Wafer Alignment.

Figure 6: Wafer Alignment

  - On each DIGITAL VIDEO IMAGE GENERATOR press 2 ONLY to display just the bottom camera display

  - Turn on the camera illumination by toggling the switch to BOTTOM OPTICS. See Figure 5: Optics.

  - Adjust the illumination knobs as needed.

  - Move each bottom camera independently along the X axis using the black knobs on the far left and right ends of the machine.

  - Move each bottom camera independently along the Y axis using the silver knobs on top of the bottom camera assembly beside the mask/wafer enclosure.

  - Focus each bottom camera independently using the BOTTOM OPTICS FOCUS located on the front side of the machine.

  - Do not set the right focus below .5 inches or it will interfere with the wafer chuck.

  - NOTE: Wafer alignment marks need to be within the viewing window for bottom side alignment. 45 mm inside, 75 mm outside.

  - Move the top objectives/camera assembly to the back of the mask plate by pressing the two control buttons on the arm.

|  | WARNING Check to ensure the objectives/camera assembly is pushed all the way back and has a clear path to move. Failure to do so will damage the objectives and their mounts. |  |
| --- | --- | --- |

- Expose Wafer

- Once the wafer is fully aligned, from the MAIN MENU press RUN SCREEN then press MOVE TO CONTACT POSITION.

- Select either CONTACT VACUUM or N2 HARD CONTACT.

- CONTACT VACUUM pulls vacuum between the mask and the wafer.

- N2 HARD CONTACT applies nitrogen under the wafer to push it against the mask.

- Make sure the objectives/camera assembly are high enough and pushed back so that they will clear the mask holder

- Press CYCLE.

- Ensure the monitors are not in the path of the lamp stage.

- Wait for UV Lamp to move over, expose the wafer, and move back over.

- Once the LEVEL button is white on the LCD screen, you are safe to pull the wafer chuck out.

- If the lamp stage does not move back to the starting position, contact NanoFab staff.

- Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 9 Aug 2016 | Cameron Anderson |  |