YES Vapor Prime Oven SOP

Scope

This document provides the procedure for operating the Yield LP III HMDS Vapor Prime Oven.

Table of Contents

1.	Scope	1

2.	Table of Contents	1

3.	Reference Documents	1

3.1	Referenced within this Document	1

3.2	External Documents	1

4.	Equipment and/or Materials	1

5.	Safety	1

6.	Setup Procedures	1

7.	Film Deposition Procedures	2

8.	Revision History	3

Figure 1, Vacuum Pump	2

Figure 2, Control Panel	3

Reference Documents

Referenced within this Document

None

External Documents

None

Equipment and/or Materials

YES Vapor Prime Oven

Wafer/Sample

HMDS

Safety

Follow all Nanofab safety procedures.

Setup Procedures

Check the status of the system: See Figure 2, Control Panel.

Vacuum Gauge: <100 Torr

Temp. ~150 C

HMDS in flask

Vacuum pump off

Turn on vacuum pump. See Figure 1, Vacuum Pump.

Press the black START button. This will vent the system for 5 minutes. See Figure 2, Control Panel.

Power Switch

Figure 1, Vacuum Pump

Film Deposition Procedures

Pull the door open and insert wafers.

Close the door. This may require a bit of force.

Press the black START button again.

NOTE:	This will start pump down and deposition process. The total time is ~30 minutes.

An alarm will sound when the run is finished.

Press the black START button again to vent the system.

Pull the door open and remove samples.

Close the door and press the black START button.

Once the vacuum gauge reads lower than 100 Torr, press the red RESET button. See Figure 2.

NOTE:	This will stop the system. Make sure this is done before the next step or it will begin a nitrogen backfill.

Turn off the vacuum pump.

HMDS Flask

Temperature

Vacuum Gauge

Start Button

Reset Button

Figure 2, Control Panel

Revision History

| Rev | Date | Originator | Description of Changes |
| --- | --- | --- | --- |
| 1 | 14 Jan 2010 | Sam Bell |  |